## DicomFull
Philips CX50 V5.x DICOM SR Export Specifications
Template | App | Label | ResultLabel | Type | internalId | laterality | hashcode | Finding Site | Concept | patientCharacteristics | Units | Target | Mode | View | Phase | Direction | Method | anatomy | branch | laterality2 | section | segment | traceMethod
Adult Echocardiography Procedure Report | Adult Echo | A Wave Amp | Length | meas | m_AWaveAmp_Distance | unilateral | 125200\59100-8\G-0394\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 59100-8, A-Wave Amplitude |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | A1 Prolapse Vol | none | meas | m_mvq_A1ProVol | unilateral | 125200\C12207-31\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-31, Mitral Valve A1 Prolapse Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | A1 Tent Vol | none | meas | m_mvq_A1TentVol | unilateral | 125200\C12207-30\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-30, Mitral Valve A1 Tenting Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | A2 Prolapse Vol | none | meas | m_mvq_A2ProVol | unilateral | 125200\C12207-33\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-33, Mitral Valve A2 Prolapse Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | A2 Tent Vol | none | meas | m_mvq_A2TentVol | unilateral | 125200\C12207-32\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-32, Mitral Valve A2 Tenting Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | A3 Prolapse Vol | none | meas | m_mvq_A3ProVol | unilateral | 125200\C12207-35\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-35, Mitral Valve A3 Prolapse Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | A3 Tent Vol | none | meas | m_mvq_A3TentVol | unilateral | 125200\C12207-34\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-34, Mitral Valve A3 Tenting Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | AI Acc Time | Time | meas | m_AI_AccelTime_DopTime | unilateral | 125200\20168-1\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AI Acc Time Slope | Doppler Slope | meas | m_AI_AccelTime_DopSlope | unilateral | 125200\20167-3\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20167-3, Acceleration Slope |  | UCUM, cm/s2, Centimeter Per Second Square |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AI Alias Vel | Unsigned Maximum Velocity | meas | m_AI_AliasVel_2DVel | unilateral | 125200\59130-5\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 59130-5, Alias velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AI Dec Slope | Doppler Slope | meas | m_AI_DecSlope_DopSlope | unilateral | 125200\20216-8\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20216-8, Deceleration Slope |  | UCUM, cm/s2, Centimeter Per Second Square |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AI Dec Slope Time | Time | meas | m_AI_DecSlope_DopTime | unilateral | 125200\20217-6\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20217-6, Deceleration Time |  | UCUM, sec, Seconds |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AI End Dias Vel | Unsigned Maximum Velocity | meas | m_AI_EndDiVel_Vel | unilateral | 125200\11653-3\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AI ERO | none | calc | c_AI_ERO | unilateral | 125200\125216\G-0367\G-038E\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Adult Echocardiography Procedure Report | Adult Echo | AI Flow Rate | none | calc | c_AI_FlowRate | unilateral | 125200\34141-2\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 34141-2, Peak Instantaneous Flow Rate |  | UCUM, ml/sec, ml/sec |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AI Fraction | none | calc | c_AI_Frac | unilateral | 125200\G-0367\G-0390\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-0390, Regurgitant Fraction |  | UCUM, %, Percent |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AI Max PG | Maximum PG | meas | m_AI_Vmax_MaxPG | unilateral | 125200\20247-3\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AI Mean PG | Mean PG | meas | m_AI_VTI_CardPGmean | unilateral | 125200\20256-4\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AI P1/2t | Pressure Half Time | meas | m_AI_PHalfT_PHalftime | unilateral | 125200\20280-4\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20280-4, Pressure Half-Time |  | UCUM, ms, Millisecond |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AI Radius | Length | meas | m_AI_Radius_Distance | unilateral | 125200\59102-4\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 59102-4, Flow Radius |  | UCUM, cm, Centimeter |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AI Vmax | Unsigned Maximum Velocity | meas | m_AI_Vmax_Vmax | unilateral | 125200\20351-3\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AI Vmean | Unsigned Time Average Peak Vel | meas | m_AI_VTI_CardTAPV | unilateral | 125200\20352-1\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20352-1, Time Averaged Mean Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AI Volume | none | calc | c_AI_Vol | unilateral | 125200\33878-0\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 33878-0, Volume Flow |  | UCUM, ml, Milliliter |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | AI VTI | Velocity Time Integral | meas | m_AI_VTI_VTI | unilateral | 125200\20354-7\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | Annul 2D Area | none | meas | m_mvq_Ann2DArea | unilateral | 125200\C12228-08\G-03A2\G-A166\T-35300\T-35313 | SRT, T-35300, Mitral Valve | SRT, G-A166, Area |  | UCUM, cm2, Square Centimeter | SRT, T-35313, Mitral Annulus | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | Annul 3D Circ | none | meas | m_mvq_Ann3DCirc | unilateral | 125200\C12207-09\P0-02242\T-35300\T-35313 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-09, Mitral Valve Annulus 3D Circumference |  | UCUM, mm, Millimeter | SRT, T-35313, Mitral Annulus | SRT, P0-02242, 3D mode
Adult Echocardiography Procedure Report | Adult Echo | Annul 3D Min Area | none | meas | m_mvq_Ann3DMinArea | unilateral | 125200\C12207-26\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-26, Mitral Valve 3D Annular Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | Annul Ellipsicity | none | meas | m_mvq_AnnEllip | unilateral | 125200\C12207-28\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-28, Mitral Valve Annulus Ellipsicity |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | Annul Height | none | meas | m_mvq_AnnHt | unilateral | 125200\C12228-08\C7470-01\G-03A2\T-35300\T-35313 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C7470-01, Height |  | UCUM, mm, Millimeter | SRT, T-35313, Mitral Annulus | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | Annul Height / Ant Post Diam | none | meas | m_mvq_AnnHPDiam | unilateral | 125200\C12207-27\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-27, Mitral Valve Annulus Height/AP Diameter Ratio |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | Ant Lat Chord | none | meas | m_mvq_ALChrd | unilateral | 125200\C12207-16\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-16, Mitral Valve Anterolateral Chord Length |  | UCUM, mm, Millimeter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | Ant Lat Post Med Diam | none | meas | m_mvq_ALPMDiam | unilateral | 125200\C12207-08\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-08, Mitral Valve Anterolateral Posteromedial Diameter |  | UCUM, mm, Millimeter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | Ant Leaflet | none | meas | m_mvq_AntLeaf | unilateral | 125200\C12207-10\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-10, Mitral Valve Anterior Leaflet Length |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | Ant Leaflet Angle | none | meas | m_mvq_ALeafAngle | unilateral | 125200\C12207-18\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-18, Mitral Valve Anterior Leaflet Angle |  | UCUM, deg, deg |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | Ant Leaflets 3D Area | none | meas | m_mvq_AntLeaf3DA | unilateral | 125200\C12207-12\C12228-08\P0-02242\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-12, Mitral Valve Anterior Leaflets 3D Area |  | UCUM, cm2, Square Centimeter |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | Ant Post Diam | none | meas | m_mvq_APDiam | unilateral | 125200\C12207-07\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-07, Mitral Valve Anterior Posterior Diameter |  | UCUM, mm, Millimeter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | Ao Arch Diam | Length | meas | m_AoArch_Diam_Distance | unilateral | 125200\18011-7\G-03A2\T-42000 | SRT, T-42000, Aorta | LN, 18011-7, Aortic Arch Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Ao Isthmus Diam | Length | meas | m_AoIsth_Diam_Distance | unilateral | 125200\18014-1\G-03A2\T-42000 | SRT, T-42000, Aorta | LN, 18014-1, Aortic Isthmus Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Ao Or – Mitr Plane Angle | none | meas | m_mvq_AoOrMitrAngle | unilateral | 125200\C12207-21\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-21, Mitral Valve Aortic Orifice to Mitral Plane Angle |  | UCUM, deg, deg |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | AoR Diam (2D) | Length | meas | m_AoR_Diam_2D_Distance | unilateral | 125200\18015-8\G-03A2\T-42000 | SRT, T-42000, Aorta | LN, 18015-8, Aortic Root Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | AoR Diam (MM) | Length | meas | m_AoR_Diam_MM_Distance | unilateral | 125200\18015-8\G-0394\T-42000 | SRT, T-42000, Aorta | LN, 18015-8, Aortic Root Diameter |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | AP2 EDA (a2DQ) | none | meas | m_2DQ_EDA_AP2_S_AR_SM | unilateral | 125200\C12228-09\C99-12228-05\G-0375\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0375, Left Ventricular Diastolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | AP2 EDA (aCMQ) | Area | meas | m_CMQ_AP2SPSM_EDA | unilateral | 125200\C12228-10\C99-12228-07\G-0375\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0375, Left Ventricular Diastolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | AP2 EDV (a2DQ) | none | meas | m_2DQ_EDV_AP2_S_MD_SM | unilateral | 125200\18026-5\C12228-09\C99-12228-05\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | AP2 EDV (aCMQ) | none | meas | m_CMQ_AP2SPMDSM_EDV | unilateral | 125200\18026-5\C12228-10\C99-12228-07\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | AP2 EF (a2DQ) | none | meas | m_2DQ_EF_AP2_S_MD_SM | unilateral | 125200\18043-0\C12228-09\C99-12228-05\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | AP2 EF (aCMQ) | none | meas | m_CMQ_AP2SPMDSM_EF | unilateral | 125200\18043-0\C12228-10\C99-12228-07\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | AP2 ESA (a2DQ) | none | meas | m_2DQ_ESA_AP2_S_AR_SM | unilateral | 125200\C12228-09\C99-12228-05\G-0374\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0374, Left Ventricular Systolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | AP2 ESA (aCMQ) | Area | meas | m_CMQ_AP2SPSM_ESA | unilateral | 125200\C12228-10\C99-12228-07\G-0374\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0374, Left Ventricular Systolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | AP2 ESV (a2DQ) | none | meas | m_2DQ_ESV_AP2_S_MD_SM | unilateral | 125200\18148-7\C12228-09\C99-12228-05\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | AP2 ESV (aCMQ) | none | meas | m_CMQ_AP2SPMDSM_ESV | unilateral | 125200\18148-7\C12228-10\C99-12228-07\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | AP2 FAC (a2DQ) | none | meas | m_2DQ_FAC_AP2_S_AR_SM | unilateral | 125200\C12228-09\C99-12228-05\G-0376\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0376, Left Ventricular Fractional Area Change |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | AP2 FAC (aCMQ) | none | meas | m_CMQ_AP2SPSM_FAC | unilateral | 125200\C12228-10\C99-12228-07\G-0376\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0376, Left Ventricular Fractional Area Change |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | AP2 Strain (aCMQ) | none | meas | m_CMQ_Str_AP2 | unilateral | 125200\C12203-66\C12228-10\C99-12228-07\G-03A2\G-A19B\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-66, Strain peak value of the Left Ventricle |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | AP3 Strain (aCMQ) | none | meas | m_CMQ_Str_AP3 | unilateral | 125200\C12203-66\C12226-04\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-66, Strain peak value of the Left Ventricle |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-04, Apical three chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | AP4 EDA (a2DQ) | none | meas | m_2DQ_EDA_AP4_S_AR_SM | unilateral | 125200\C12228-09\C99-12228-05\G-0375\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0375, Left Ventricular Diastolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | AP4 EDA (aCMQ) | Area | meas | m_CMQ_AP4SPSM_EDA | unilateral | 125200\C12228-10\C99-12228-07\G-0375\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0375, Left Ventricular Diastolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | AP4 EDV (a2DQ) | none | meas | m_2DQ_EDV_AP4_S_MD_SM | unilateral | 125200\18026-5\C12228-09\C99-12228-05\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | AP4 EDV (aCMQ) | none | meas | m_CMQ_AP4SPMDSM_EDV | unilateral | 125200\18026-5\C12228-10\C99-12228-07\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | AP4 EF (a2DQ) | none | meas | m_2DQ_EF_AP4_S_MD_SM | unilateral | 125200\18043-0\C12228-09\C99-12228-05\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | AP4 EF (aCMQ) | none | meas | m_CMQ_AP4SPMDSM_EF | unilateral | 125200\18043-0\C12228-10\C99-12228-07\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | AP4 ESA (a2DQ) | none | meas | m_2DQ_ESA_AP4_S_AR_SM | unilateral | 125200\C12228-09\C99-12228-05\G-0374\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0374, Left Ventricular Systolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | AP4 ESA (aCMQ) | Area | meas | m_CMQ_AP4SPSM_ESA | unilateral | 125200\C12228-10\C99-12228-07\G-0374\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0374, Left Ventricular Systolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | AP4 ESV (a2DQ) | none | meas | m_2DQ_ESV_AP4_S_MD_SM | unilateral | 125200\18148-7\C12228-09\C99-12228-05\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | AP4 ESV (aCMQ) | none | meas | m_CMQ_AP4SPMDSM_ESV | unilateral | 125200\18148-7\C12228-10\C99-12228-07\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | AP4 FAC (a2DQ) | none | meas | m_2DQ_FAC_AP4_S_AR_SM | unilateral | 125200\C12228-09\C99-12228-05\G-0376\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0376, Left Ventricular Fractional Area Change |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | AP4 FAC (aCMQ) | none | meas | m_CMQ_AP4SPSM_FAC | unilateral | 125200\C12228-10\C99-12228-07\G-0376\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0376, Left Ventricular Fractional Area Change |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | AP4 Strain (aCMQ) | none | meas | m_CMQ_Str_AP4 | unilateral | 125200\C12203-66\C99-12243-01\G-03A2\G-A19C\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-66, Strain peak value of the Left Ventricle |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C99-12243-01, Lagrangian strain for the longitudinal regions
Adult Echocardiography Procedure Report | Adult Echo | Asc Ao Diam | Length | meas | m_AscAorta_Diam_Distance | unilateral | 125200\18012-5\G-03A2\T-42000 | SRT, T-42000, Aorta | LN, 18012-5, Ascending Aortic Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | Associated Morphology | none | meas | #AssocMorph | n/a | 125200\G-C504 |  | SRT, G-C504, Associated Morphology
Adult Echocardiography Procedure Report | Adult Echo | AV Acc Time | Time | meas | m_AV_AccelTime_DopTime | unilateral | 125200\20168-1\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | AV Acc Time Slope | Doppler Slope | meas | m_AV_AccelTime_DopSlope | unilateral | 125200\20167-3\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20167-3, Acceleration Slope |  | UCUM, cm/s2, Centimeter Per Second Square |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | AV Area | Area | meas | m_AV_Area_Area | unilateral | 125200\125220\G-038E\G-03A2\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125220, Planimetry
Adult Echocardiography Procedure Report | Adult Echo | AV Cusp Sep | Length | meas | m_AV_Cusp_Sep_Distance | unilateral | 125200\17996-0\G-0394\T-35400 | SRT, T-35400, Aortic Valve | LN, 17996-0, Aortic Valve Cusp Separation |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | AV Dec Time | Time | meas | m_AV_DecelTime_DopTime | unilateral | 125200\20217-6\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20217-6, Deceleration Time |  | UCUM, sec, Seconds |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | AV Max PG | Maximum PG | meas | m_AV_Vmax_MaxPG | unilateral | 125200\20247-3\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | AV Mean PG | Mean PG | meas | m_AV_VTI_CardPGmean | unilateral | 125200\20256-4\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | AV R-R | Time | meas | m_HR_AV_HeartRate | unilateral | 125200\8867-4\T-35400 | SRT, T-35400, Aortic Valve | LN, 8867-4, Heart Rate |  | UCUM, sec, Seconds
Adult Echocardiography Procedure Report | Adult Echo | AV Vmax | Unsigned Maximum Velocity | meas | m_AV_Vmax_Vmax | unilateral | 125200\20351-3\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | AV Vmean | Unsigned Time Average Peak Vel | meas | m_AV_VTI_CardTAPV | unilateral | 125200\20352-1\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20352-1, Time Averaged Mean Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | AV VTI | Velocity Time Integral | meas | m_AV_VTI_VTI | unilateral | 125200\20354-7\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | AVA (Vmax) | none | calc | c_AVA_Vmax | unilateral | 125200\125214\G-038E\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | [5 empty] | DCM, 125214, Continuity Equation by Peak Velocity
Adult Echocardiography Procedure Report | Adult Echo | AVA (VTI) | none | calc | c_AVA_VTI | unilateral | 125200\125215\G-038E\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | [5 empty] | DCM, 125215, Continuity Equation by Velocity Time Integral
Adult Echocardiography Procedure Report | Adult Echo | B-C Slope | MMode Slope | meas | m_BC_Slope_MMSlope | unilateral | 125200\59126-3\G-0394\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 59126-3, B-C Slope |  | UCUM, cm/s, Centimeter Per Second |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | Bi-Plane EDV (a2DQ) | none | meas | m_2DQ_EDV_AP_BP_MD_SM | unilateral | 125200\18026-5\C12228-09\C99-12228-05\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | Bi-Plane EDV (aCMQ) | none | meas | m_CMQ_BP_MD_EDV | unilateral | 125200\18026-5\C12228-10\C99-12228-07\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | Bi-Plane EF (a2DQ) | none | meas | m_2DQ_EF_AP_BP_MD_SM | unilateral | 125200\18043-0\C12228-09\C99-12228-05\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | Bi-Plane EF (aCMQ) | none | meas | m_CMQ_BP_MD_EF | unilateral | 125200\18043-0\C12228-10\C99-12228-07\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | Bi-Plane ESV (a2DQ) | none | meas | m_2DQ_ESV_AP_BP_MD_SM | unilateral | 125200\18148-7\C12228-09\C99-12228-05\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | Bi-Plane ESV (aCMQ) | none | meas | m_CMQ_BP_MD_ESV | unilateral | 125200\18148-7\C12228-10\C99-12228-07\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | Cardiac Wall Motion | none | meas | #CardWallMotion | n/a | 125200\F-32050 |  | SRT, F-32050, Cardiac Wall Motion
Adult Echocardiography Procedure Report | Adult Echo | CI (2D-Cubed) | none | calc | c_CI_Cubed | unilateral | 125200\125206\F-32110\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | CI (2D-Teich) | none | calc | c_CI_Teich | unilateral | 125200\125209\F-32110\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | CI (A/L) | none | calc | c_CI_SPE | unilateral | 125200\125226\F-32110\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | CI (MM-Cubed) | none | calc | c_CI_Cubed_MM | unilateral | 125200\125206\F-32110\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | CI (MM-Teich) | none | calc | c_CI_Teich_MM | unilateral | 125200\125209\F-32110\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | CI(MOD-bp) | none | calc | c_CI_MOD_BP | unilateral | 125200\125207\F-32110\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Adult Echocardiography Procedure Report | Adult Echo | CI(MOD-sp2) | none | calc | c_CI_MOD_SP2 | unilateral | 125200\125208\F-32110\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | CI(MOD-sp4) | none | calc | c_CI_MOD_SP4 | unilateral | 125200\125208\F-32110\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | CO (2D-Cubed) | none | calc | c_CO_Cubed | unilateral | 125200\125206\F-32100\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | CO (2D-Teich) | none | calc | c_CO_Teich | unilateral | 125200\125209\F-32100\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | CO (A/L) | none | calc | c_CO_SPE | unilateral | 125200\125226\F-32100\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | CO (LVOT) | none | calc | c_CO_LVOT | unilateral | 125200\F-32100\T-32600\T-32650 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute | SRT, T-32650, Left Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | CO (MM-Cubed) | none | calc | c_CO_Cubed_MM | unilateral | 125200\125206\F-32100\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | CO (MM-Teich) | none | calc | c_CO_Teich_MM | unilateral | 125200\125209\F-32100\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | CO (MV) | none | calc | c_CO_MV | unilateral | 125200\F-32100\T-35300 | SRT, T-35300, Mitral Valve | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute
Adult Echocardiography Procedure Report | Adult Echo | CO (RVOT) | none | calc | c_CO_RVOT | unilateral | 125200\F-32100\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute | SRT, T-32550, Right Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | CO (TV) | none | calc | c_CO_TV | unilateral | 125200\F-32100\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute
Adult Echocardiography Procedure Report | Adult Echo | CO(MOD-bp) | none | calc | c_CO_MOD_BP | unilateral | 125200\125207\F-32100\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Adult Echocardiography Procedure Report | Adult Echo | CO(MOD-sp2) | none | calc | c_CO_MOD_SP2 | unilateral | 125200\125208\F-32100\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | CO(MOD-sp4) | none | calc | c_CO_MOD_SP4 | unilateral | 125200\125208\F-32100\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | Coapt 2D | none | meas | m_mvq_Coapt2D | unilateral | 125200\C12207-22\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-22, Mitral Valve Coaptation 2D Length |  | UCUM, mm, Millimeter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | Desc Ao Diam | Length | meas | m_DescAorta_Diam_Distance | unilateral | 125200\18013-3\G-03A2\T-42000 | SRT, T-42000, Aorta | LN, 18013-3, Descending Aortic Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | E/Lat E` | none | calc | c_E_LatE | unilateral | 125200\G-037B\G-0392\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037B, Ratio of MV Peak Velocity to LV Peak Tissue Velocity E-Wave |  | UCUM, 1, no units | SRT, G-0392, Lateral Mitral Annulus
Adult Echocardiography Procedure Report | Adult Echo | E/Med E` | none | calc | c_E_MedE | unilateral | 125200\G-037B\G-0391\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037B, Ratio of MV Peak Velocity to LV Peak Tissue Velocity E-Wave |  | UCUM, 1, no units | SRT, G-0391, Medial Mitral Annulus
Adult Echocardiography Procedure Report | Adult Echo | E`/A` Lateral | none | calc | c_Lat_E_A | unilateral | 125200\59129-7\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59129-7, Left Ventricle E to A Tissue Velocity Ratio |  | UCUM, 1, no units | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | E`/A` Medial | none | calc | c_Med_E_A | unilateral | 125200\59129-7\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59129-7, Left Ventricle E to A Tissue Velocity Ratio |  | UCUM, 1, no units | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | EDV (2D-Cubed) | none | calc | c_EDV_Cubed | unilateral | 125200\125206\18026-5\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | EDV (2D-Teich) | none | calc | c_EDV_Teich | unilateral | 125200\125209\18026-5\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | EDV (A/L) | none | calc | c_EDV_SPE | unilateral | 125200\125226\18026-5\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | EDV (MM-Cubed) | none | calc | c_EDV_Cubed_MM | unilateral | 125200\125206\18026-5\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | EDV (MM-Teich) | none | calc | c_EDV_Teich_MM | unilateral | 125200\125209\18026-5\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | EDV 3D-VOL | none | meas | m_EDV_3D_VOL | unilateral | 125200\18026-5\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  | UCUM, ml, Milliliter |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | EDV BP | none | meas | m_EDV_BP_TEMPL | unilateral | 125200\18026-5\C12228-06\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  | UCUM, ml, Milliliter |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-06, QLab 3DQ
Adult Echocardiography Procedure Report | Adult Echo | EDV(MOD-bp) | none | calc | c_EDV_MOD_BP | unilateral | 125200\125207\18026-5\F-32011\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole |  | DCM, 125207, Method of Disks, Biplane
Adult Echocardiography Procedure Report | Adult Echo | EDV(MOD-sp2) | MOD Volume | meas | m_EDV_MOD_SP2 | unilateral | 125200\125208\18026-5\F-32011\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | EDV(MOD-sp4) | MOD Volume | meas | m_EDV_MOD_SP4 | unilateral | 125200\125208\18026-5\F-32011\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | EF (2D-Cubed) | none | calc | c_EF_Cubed | unilateral | 125200\125206\18043-0\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | EF (2D-Teich) | none | calc | c_EF_Teich | unilateral | 125200\125209\18043-0\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | EF (A/L) | none | calc | c_EF_SPE | unilateral | 125200\125226\18043-0\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | EF (Dumesnil) | none | calc | c_EF_Dumesnil | unilateral | 125200\C3467-04\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C3467-04, Left Ventricular Ejection Fraction by Dumesnil 1995 |  | UCUM, %, Percent
Adult Echocardiography Procedure Report | Adult Echo | EF (MM-Cubed) | none | calc | c_EF_Cubed_MM | unilateral | 125200\125206\18043-0\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  | UCUM, %, Percent |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | EF (MM-Teich) | none | calc | c_EF_Teich_MM | unilateral | 125200\125209\18043-0\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  | UCUM, %, Percent |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | EF 3D-VOL | none | meas | m_EF_3D_VOL | unilateral | 125200\18043-0\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  | UCUM, %, Percent |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | EF BP | none | meas | m_EF_BP_TEMPL | unilateral | 125200\18043-0\C12228-06\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  | UCUM, %, Percent |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-06, QLab 3DQ
Adult Echocardiography Procedure Report | Adult Echo | EF(MOD-bp) | none | calc | c_EF_MOD_BP | unilateral | 125200\125207\18043-0\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Adult Echocardiography Procedure Report | Adult Echo | EF(MOD-sp2) | none | calc | c_EF_MOD_SP2 | unilateral | 125200\125208\18043-0\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | EF(MOD-sp4) | none | calc | c_EF_MOD_SP4 | unilateral | 125200\125208\18043-0\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | LN, 18043-0, Left Ventricular Ejection Fraction |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | ESV (2D-Cubed) | none | calc | c_ESV_Cubed | unilateral | 125200\125206\18148-7\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | ESV (2D-Teich) | none | calc | c_ESV_Teich | unilateral | 125200\125209\18148-7\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | ESV (A/L) | none | calc | c_ESV_SPE | unilateral | 125200\125226\18148-7\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | ESV (MM-Cubed) | none | calc | c_ESV_Cubed_MM | unilateral | 125200\125206\18148-7\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | ESV (MM-Teich) | none | calc | c_ESV_Teich_MM | unilateral | 125200\125209\18148-7\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | ESV 3D-VOL | none | meas | m_ESV_3D_VOL | unilateral | 125200\18148-7\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  | UCUM, ml, Milliliter |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | ESV BP | none | meas | m_ESV_BP_TEMPL | unilateral | 125200\18148-7\C12228-06\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  | UCUM, ml, Milliliter |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-06, QLab 3DQ
Adult Echocardiography Procedure Report | Adult Echo | ESV(MOD-bp) | none | calc | c_ESV_MOD_BP | unilateral | 125200\125207\18148-7\G-03A2\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  | SRT, R-FAB5B, End Systole |  | DCM, 125207, Method of Disks, Biplane
Adult Echocardiography Procedure Report | Adult Echo | ESV(MOD-sp2) | MOD Volume | meas | m_ESV_MOD_SP2 | unilateral | 125200\125208\18148-7\G-03A2\G-A19B\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | ESV(MOD-sp4) | MOD Volume | meas | m_ESV_MOD_SP4 | unilateral | 125200\125208\18148-7\G-03A2\G-A19C\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | FS (2D-Cubed) | none | calc | c_FS_Cubed | unilateral | 125200\125206\18051-3\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18051-3, Left Ventricular Fractional Shortening |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | FS (2D-Teich) | none | calc | c_FS_Teich | unilateral | 125200\125209\18051-3\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18051-3, Left Ventricular Fractional Shortening |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | FS (MM-Cubed) | none | calc | c_FS_Cubed_MM | unilateral | 125200\125206\18051-3\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18051-3, Left Ventricular Fractional Shortening |  | UCUM, %, Percent |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | FS (MM-Teich) | none | calc | c_FS_Teich_MM | unilateral | 125200\125209\18051-3\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18051-3, Left Ventricular Fractional Shortening |  | UCUM, %, Percent |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | Global C Strain (aCMQ) | none | meas | m_CMQ_GStr_SAX | unilateral | 125200\C12203-66\C12226-08\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-66, Strain peak value of the Left Ventricle |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-08, Combination of SAX A, SAX M and SAX B views |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | Global L Strain (aCMQ) | none | meas | m_CMQ_GStr_AP | unilateral | 125200\C12203-66\C12226-02\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-66, Strain peak value of the Left Ventricle |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-02, Combination of AP2, AP3 and AP4 views |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | Hep. A Revs Dur Time | Time | meas | m_HepARevsDur_DopTime | unilateral | 125200\C12216-01\T-48720 | SRT, T-48720, Hepatic Vein | 99PMSBLUS, C12216-01, Hepatic Vein A-Wave Duration |  | UCUM, sec, Seconds
Adult Echocardiography Procedure Report | Adult Echo | Hep. A Revs Vel | Unsigned Maximum Velocity | meas | m_HepARevsVel_Vel | unilateral | 125200\29474-4\T-48720 | SRT, T-48720, Hepatic Vein | LN, 29474-4, Hepatic Vein Atrial Contraction Reversal Peak Velocity |  | UCUM, cm/s, Centimeter Per Second
Adult Echocardiography Procedure Report | Adult Echo | Hepatic Dias Vel | Unsigned Maximum Velocity | meas | m_HepDiasVel_Vel | unilateral | 125200\29472-8\T-48720 | SRT, T-48720, Hepatic Vein | LN, 29472-8, Hepatic Vein Diastolic Peak Velocity |  | UCUM, cm/s, Centimeter Per Second
Adult Echocardiography Procedure Report | Adult Echo | Hepatic S/D | none | calc | c_HepSD | unilateral | 125200\29473-6\T-48720 | SRT, T-48720, Hepatic Vein | LN, 29473-6, Hepatic Vein Systolic to Diastolic Ratio |  | UCUM, 1, no units
Adult Echocardiography Procedure Report | Adult Echo | Hepatic Sys Vel | Unsigned Maximum Velocity | meas | m_HepSysVel_Vel | unilateral | 125200\29471-0\T-48720 | SRT, T-48720, Hepatic Vein | LN, 29471-0, Hepatic Vein Systolic Peak Velocity |  | UCUM, cm/s, Centimeter Per Second
Adult Echocardiography Procedure Report | Adult Echo | HR LV | none | calc | c_MM_DIM_HR | unilateral | 125200\8867-4\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, Beats Per Minute |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | IVC Diameter | Length | meas | m_IVCDiam_2D_Distance | unilateral | 125200\18006-7\G-03A2\T-48600 | SRT, T-48600, Vena Cava | LN, 18006-7, Inferior Vena Cava Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | IVCT Time | Time | meas | m_IVCT_DopTime | unilateral | 125200\G-037E\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037E, Left Ventricular Isovolumic Contraction Time |  | UCUM, sec, Seconds
Adult Echocardiography Procedure Report | Adult Echo | IVRT Time | Time | meas | m_IVRT_DopTime | unilateral | 125200\18071-1\T-32600 | SRT, T-32600, Left Ventricle | LN, 18071-1, Left Ventricular Isovolumic Relaxation Time |  | UCUM, sec, Seconds
Adult Echocardiography Procedure Report | Adult Echo | IVS % (2D) | none | calc | c_IVS | unilateral | 125200\18054-7\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18054-7, Interventricular Septum % Thickening |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | IVS % (MM) | none | calc | c_IVS_MM | unilateral | 125200\18054-7\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18054-7, Interventricular Septum % Thickening |  | UCUM, %, Percent |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | IVS/LVPW (2D) | none | calc | c_IVS_LVPW | unilateral | 125200\18155-2\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18155-2, Interventricular Septum to Posterior Wall Thickness Ratio |  | UCUM, 1, no units |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | IVS/LVPW (MM) | none | calc | c_IVS_LVPW_MM | unilateral | 125200\18155-2\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18155-2, Interventricular Septum to Posterior Wall Thickness Ratio |  | UCUM, 1, no units |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | IVSd (2D) | Length | meas | m_IVSd_2D_Distance | unilateral | 125200\18154-5\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18154-5, Interventricular Septum Diastolic Thickness |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | IVSd (MM) | Length | meas | m_IVSd_MM_Distance | unilateral | 125200\18154-5\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18154-5, Interventricular Septum Diastolic Thickness |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | IVSs (2D) | Length | meas | m_IVSs_2D_Distance | unilateral | 125200\18158-6\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18158-6, Interventricular Septum Systolic Thickness |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | IVSs (MM) | Length | meas | m_IVSs_MM_Distance | unilateral | 125200\18158-6\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18158-6, Interventricular Septum Systolic Thickness |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | LA Dimen (2D) | Length | meas | m_LA_Dimen_2D_Distance | unilateral | 125200\29469-4\G-03A2\T-32300 | SRT, T-32300, Left Atrium | LN, 29469-4, Left Atrium Antero-posterior Systolic Dimension |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LA Dimen (MM) | Length | meas | m_LA_Dimen_MM_Distance | unilateral | 125200\29469-4\G-0394\T-32300 | SRT, T-32300, Left Atrium | LN, 29469-4, Left Atrium Antero-posterior Systolic Dimension |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | LA/Ao (2D) | none | calc | c_LA_Ao_2D | unilateral | 125200\17985-3\G-03A2\T-32300 | SRT, T-32300, Left Atrium | LN, 17985-3, Left Atrium to Aortic Root Ratio |  | UCUM, 1, no units |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LA/Ao (MM) | none | calc | c_LA_Ao_MM | unilateral | 125200\17985-3\G-0394\T-32300 | SRT, T-32300, Left Atrium | LN, 17985-3, Left Atrium to Aortic Root Ratio |  | UCUM, 1, no units |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | Lat A` Area VTI | Velocity Time Integral | meas | m_AreaUnderLatPeakA_VTI | unilateral | 125200\59125-5\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59125-5, LV VTI A wave US |  | UCUM, cm, Centimeter | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Lat A` Vel | Unsigned Maximum Velocity | meas | m_Lat_PeakAVel_Vel | unilateral | 125200\G-037C\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037C, LV Peak Diastolic Tissue Velocity During Atrial Systole |  | UCUM, cm/s, Centimeter Per Second | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Lat Acc Time | Time | meas | m_TDILatAccelTime_DopTime | unilateral | 125200\20168-1\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Lat Dec Time | Time | meas | m_TDILatDecelTime_DopTime | unilateral | 125200\20217-6\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 20217-6, Deceleration Time |  | UCUM, sec, Seconds | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Lat E` Area VTI | Velocity Time Integral | meas | m_AreaUnderLatPeakE_VTI | unilateral | 125200\59124-8\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59124-8, LV VTI E wave US |  | UCUM, cm, Centimeter | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Lat E` Vel | Unsigned Maximum Velocity | meas | m_Lat_PeakEVel_Vel | unilateral | 125200\G-037A\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037A, Left Ventricular Peak Early Diastolic Tissue Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Lat IVCT Time | Time | meas | m_TDI_Lat_IVCT_DopTime | unilateral | 125200\G-037E\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037E, Left Ventricular Isovolumic Contraction Time |  | UCUM, sec, Seconds | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Lat IVRT Time | Time | meas | m_TDI_Lat_IVRT_DopTime | unilateral | 125200\18071-1\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 18071-1, Left Ventricular Isovolumic Relaxation Time |  | UCUM, sec, Seconds | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Lat S Vel | Unsigned Maximum Velocity | meas | m_Lat_PeakSVel_Vel | unilateral | 125200\G-037D\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037D, Left Ventricular Peak Systolic Tissue Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Late Dias Slope | MMode Slope | meas | m_LateDiasSlope_MMSlope | unilateral | 125200\C12209-01\G-0394\T-35200 | SRT, T-35200, Pulmonic Valve | 99PMSBLUS, C12209-01, Late Diastolic Slope |  | UCUM, cm/s, Centimeter Per Second |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | Leaflets 3D Area | none | meas | m_mvq_Leaf3DArea | unilateral | 125200\C12207-44\C12228-08\P0-02242\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-44, Mitral Valve Leaflets 3D Area |  | UCUM, cm2, Square Centimeter |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | LPA Diam | Length | meas | m_LPA_Diam_Distance | unilateral | 125200\18019-0\G-03A2\T-44000 | SRT, T-44000, Pulmonary artery | LN, 18019-0, Left Pulmonary Artery Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LV Dp/dt | Doppler Dp/dt | meas | m_LV_dpdt | unilateral | 125200\18035-6\T-35300 | SRT, T-35300, Mitral Valve | LN, 18035-6, Mitral Regurgitation dP/dt derived from Mitral Reg velocity |  | UCUM, mm[Hg]/s, mmHg/s
Adult Echocardiography Procedure Report | Adult Echo | LV EDA (PSAX) | Area | meas | m_LVEDA_2D_Area | unilateral | 125200\F-32011\G-0375\G-039B\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0375, Left Ventricular Diastolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-039B, Parasternal short axis at the Papillary Muscle level | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | LV ESA (PSAX) | Area | meas | m_LVESA_2D_Area | unilateral | 125200\G-0374\G-039B\G-03A2\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0374, Left Ventricular Systolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-039B, Parasternal short axis at the Papillary Muscle level | SRT, R-FAB5B, End Systole
Adult Echocardiography Procedure Report | Adult Echo | LV ET Time | Time | meas | m_LV_Eject_Tm_MMTime | unilateral | 125200\20222-6\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 20222-6, Ejection Time |  | UCUM, sec, Seconds |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | LV Excursion - Avg | none | meas | m_3DQA_LVExcMean | unilateral | 125200\C12203-27\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-27, The Avg Inward Motion of the Myocardium Across the LV Wall |  | UCUM, mm, Millimeter |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | LV Excursion - Max | none | meas | m_3DQA_LVExcMax | unilateral | 125200\C12203-30\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-30, The Max Inward Motion of the Myocardium Across the LV Wall |  | UCUM, mm, Millimeter |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | LV Excursion - Min | none | meas | m_3DQA_LVExcMin | unilateral | 125200\C12203-29\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-29, The Min Inward Motion of the Myocardium Across the LV Wall |  | UCUM, mm, Millimeter |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | LV Excursion - SD | none | meas | m_3DQA_LVExcSD | unilateral | 125200\C12203-28\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-28, The SD Inward Motion of the Myocardium Across the LV Wall |  | UCUM, mm, Millimeter |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | LV FAC (PSAX) | none | calc | c_LV_FAC | unilateral | 125200\G-0376\G-039B\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0376, Left Ventricular Fractional Area Change |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-039B, Parasternal short axis at the Papillary Muscle level
Adult Echocardiography Procedure Report | Adult Echo | LV Mass (A/L) | none | calc | c_LVMass_A_L | unilateral | 125200\18087-7\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18087-7, Left Ventricle Mass |  | UCUM, g, Gram |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LV Mass (Cubed) | none | calc | c_LV_Mass | unilateral | 125200\125206\18087-7\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18087-7, Left Ventricle Mass |  | UCUM, g, Gram |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | LV Mass Index (A/L) | none | calc | c_LVMass_A_LBSA | unilateral | 125200\C12203-01\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-01, Left Ventricle Mass Index |  | UCUM, g/m2, g/m2
Adult Echocardiography Procedure Report | Adult Echo | LV Mass Index(Cubed) | none | calc | c_LV_Mass_BSA | unilateral | 125200\125206\C12203-01\G-0394\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-01, Left Ventricle Mass Index |  | UCUM, g/m2, g/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | LV PEP Time | Time | meas | m_LV_PreEject_Tm_MMTime | unilateral | 125200\59085-1\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59085-1, Pre-Ejection Period |  | UCUM, sec, Seconds |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | LV PEP/ET | none | calc | c_LV_PEP_ET | unilateral | 125200\C12203-04\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-04, PEP/ET |  | UCUM, 1, no units
Adult Echocardiography Procedure Report | Adult Echo | LV R-R | Time | meas | m_MM_DIM_R_R_INT | unilateral | 125200\8867-4\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 8867-4, Heart Rate |  | UCUM, sec, Seconds |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | LV Wall Motion Score Index | none | meas | #LVWallMotionScore | n/a | 125200\125202 |  | SRT, 125202, LV Wall Motion Score Index
Adult Echocardiography Procedure Report | Adult Echo | LV Wall Motion Segmental Findings | none | meas | #LVWallMotionSeg | n/a | 125200\18118-0 |  | LN, 18118-0, LV Wall Motion Segmental Findings
Adult Echocardiography Procedure Report | Adult Echo | LVAd (A/L) | Area | meas | m_LVAd_Apical | unilateral | 125200\125226\F-32011\G-0375\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0375, Left Ventricular Diastolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | LVAd ap2 | Area | meas | m_LVAD_AP2 | unilateral | 125200\125208\F-32011\G-0375\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0375, Left Ventricular Diastolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | LVAd ap4 | Area | meas | m_LVAD_AP4 | unilateral | 125200\125208\F-32011\G-0375\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0375, Left Ventricular Diastolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | LVAd Sax Endo Area | Area | meas | m_LV_AdSaxEndo_Area | unilateral | 125200\G-0375\G-039B\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0375, Left Ventricular Diastolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-039B, Parasternal short axis at the Papillary Muscle level
Adult Echocardiography Procedure Report | Adult Echo | LVAd Sax Epi Area | Area | meas | m_LV_AdSaxEpi_Area | unilateral | 125200\G-0379\G-039B\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0379, Left Ventricle Epicardial Diastolic Area, psax pap view |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-039B, Parasternal short axis at the Papillary Muscle level
Adult Echocardiography Procedure Report | Adult Echo | LVAs (A/L) | Area | meas | m_LVAs_Apical | unilateral | 125200\125226\G-0374\G-03A2\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0374, Left Ventricular Systolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode |  | SRT, R-FAB5B, End Systole |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | LVAs ap2 | Area | meas | m_LVAS_AP2 | unilateral | 125200\125208\G-0374\G-03A2\G-A19B\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0374, Left Ventricular Systolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | LVAs ap4 | Area | meas | m_LVAS_AP4 | unilateral | 125200\125208\G-0374\G-03A2\G-A19C\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0374, Left Ventricular Systolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | LVIDd (2D) | Length | meas | m_LVIDd_2D_Distance | unilateral | 125200\29436-3\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 29436-3, Left Ventricle Internal End Diastolic Dimension |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LVIDd (MM) | Length | meas | m_LVIDd_MM_Distance | unilateral | 125200\29436-3\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 29436-3, Left Ventricle Internal End Diastolic Dimension |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | LVIDs (2D) | Length | meas | m_LVIDs_2D_Distance | unilateral | 125200\29438-9\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 29438-9, Left Ventricle Internal Systolic Dimension |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LVIDs (MM) | Length | meas | m_LVIDs_MM_Distance | unilateral | 125200\29438-9\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 29438-9, Left Ventricle Internal Systolic Dimension |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | LVLd (A/L) | Length | meas | m_LVLd_Apical | unilateral | 125200\125226\18077-8\F-32011\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18077-8, Left Ventricle diastolic major axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | LVLd Apical | Length | meas | m_LV_LdApical_Distance | unilateral | 125200\18077-8\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18077-8, Left Ventricle diastolic major axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LVLs (A/L) | Length | meas | m_LVLs_Apical | unilateral | 125200\125226\18076-0\G-03A2\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 18076-0, Left Ventricle systolic major axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode |  | SRT, R-FAB5B, End Systole |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | LVOT Acc Time | Time | meas | m_LVOT_AccelTime_DopTime | unilateral | 125200\20168-1\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | SRT, T-32650, Left Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | LVOT Acc Time Slope | Doppler Slope | meas | m_LVOT_AccelTime_DopSlope | unilateral | 125200\20167-3\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20167-3, Acceleration Slope |  | UCUM, cm/s2, Centimeter Per Second Square | SRT, T-32650, Left Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | LVOT Area | none | calc | c_LVOT_Area | unilateral | 125200\G-038E\G-03A2\T-32600\T-32650 | SRT, T-32600, Left Ventricle | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | SRT, T-32650, Left Ventricle Outflow Tract | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LVOT Diam | Length | meas | m_LVOT_Diam_Distance | unilateral | 125200\G-038F\T-32600\T-32650 | SRT, T-32600, Left Ventricle | SRT, G-038F, Cardiovascular Orifice Diameter |  | UCUM, cm, Centimeter | SRT, T-32650, Left Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | LVOT Max PG | Maximum PG | meas | m_LVOT_Vmax_MaxPG | unilateral | 125200\20247-3\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury | SRT, T-32650, Left Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | LVOT Mean PG | Mean PG | meas | m_LVOT_VTI_CardPGmean | unilateral | 125200\20256-4\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury | SRT, T-32650, Left Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | LVOT Vmax | Unsigned Maximum Velocity | meas | m_LVOT_Vmax_Vmax | unilateral | 125200\20351-3\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, T-32650, Left Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | LVOT Vmean | Unsigned Time Average Peak Vel | meas | m_LVOT_VTI_CardTAPV | unilateral | 125200\20352-1\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20352-1, Time Averaged Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, T-32650, Left Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | LVOT VTI | Velocity Time Integral | meas | m_LVOT_VTI_VTI | unilateral | 125200\20354-7\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | SRT, T-32650, Left Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | LVPW % (2D) | none | calc | c_LVPW | unilateral | 125200\18053-9\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18053-9, Left Ventricle Posterior Wall % Thickening |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LVPW % (MM) | none | calc | c_LVPW_MM | unilateral | 125200\18053-9\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18053-9, Left Ventricle Posterior Wall % Thickening |  | UCUM, %, Percent |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | LVPWd (2D) | Length | meas | m_LVPWd_2D_Distance | unilateral | 125200\18152-9\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18152-9, Left Ventricle Posterior Wall Diastolic Thickness |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LVPWd (MM) | Length | meas | m_LVPWd_MM_Distance | unilateral | 125200\18152-9\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18152-9, Left Ventricle Posterior Wall Diastolic Thickness |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | LVPWs (2D) | Length | meas | m_LVPWs_2D_Distance | unilateral | 125200\18156-0\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18156-0, Left Ventricle Posterior Wall Systolic Thickness |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | LVPWs (MM) | Length | meas | m_LVPWs_MM_Distance | unilateral | 125200\18156-0\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18156-0, Left Ventricle Posterior Wall Systolic Thickness |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | Med A` Area VTI | Velocity Time Integral | meas | m_AreaUnderMedPeakA_VTI | unilateral | 125200\59125-5\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59125-5, LV VTI A wave US |  | UCUM, cm, Centimeter | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Med A` Vel | Unsigned Maximum Velocity | meas | m_Med_PeakAVel_Vel | unilateral | 125200\G-037C\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037C, LV Peak Diastolic Tissue Velocity During Atrial Systole |  | UCUM, cm/s, Centimeter Per Second | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Med Acc Time | Time | meas | m_TDIMedAccelTime_DopTime | unilateral | 125200\20168-1\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Med Dec Time | Time | meas | m_TDIMedDecelTime_DopTime | unilateral | 125200\20217-6\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 20217-6, Deceleration Time |  | UCUM, sec, Seconds | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Med E` Area VTI | Velocity Time Integral | meas | m_AreaUnderMedPeakE_VTI | unilateral | 125200\59124-8\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59124-8, LV VTI E wave US |  | UCUM, cm, Centimeter | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Med E` Vel | Unsigned Maximum Velocity | meas | m_Med_PeakEVel_Vel | unilateral | 125200\G-037A\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037A, Left Ventricular Peak Early Diastolic Tissue Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Med IVCT Time | Time | meas | m_TDI_Med_IVCT_DopTime | unilateral | 125200\G-037E\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037E, Left Ventricular Isovolumic Contraction Time |  | UCUM, sec, Seconds | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Med IVRT Time | Time | meas | m_TDI_Med_IVRT_DopTime | unilateral | 125200\18071-1\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 18071-1, Left Ventricular Isovolumic Relaxation Time |  | UCUM, sec, Seconds | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Med S Vel | Unsigned Maximum Velocity | meas | m_Med_PeakSVel_Vel | unilateral | 125200\G-037D\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037D, Left Ventricular Peak Systolic Tissue Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | MPA Diam | Length | meas | m_MPA_Diam_Distance | unilateral | 125200\18020-8\G-03A2\T-44000 | SRT, T-44000, Pulmonary artery | LN, 18020-8, Main Pulmonary Artery Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | MR Alias Vel | Unsigned Maximum Velocity | meas | m_MR_AliasVel_2DVel | unilateral | 125200\59130-5\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 59130-5, Alias velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | MR ERO | none | calc | c_MR_ERO | unilateral | 125200\125216\G-0367\G-038E\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Adult Echocardiography Procedure Report | Adult Echo | MR Flow Rate | none | calc | c_MR_FlowRate | unilateral | 125200\34141-2\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 34141-2, Peak Instantaneous Flow Rate |  | UCUM, ml/sec, ml/sec |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | MR Fraction | none | calc | c_MR_Frac | unilateral | 125200\G-0367\G-0390\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-0390, Regurgitant Fraction |  | UCUM, %, Percent |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | MR Max PG | Maximum PG | meas | m_MR_Vmax_MaxPG | unilateral | 125200\20247-3\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | MR Mean PG | Mean PG | meas | m_MR_VTI_CardPGmean | unilateral | 125200\20256-4\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | MR Radius | Length | meas | m_MR_Radius_Distance | unilateral | 125200\59102-4\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 59102-4, Flow Radius |  | UCUM, cm, Centimeter |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | MR Vmax | Unsigned Maximum Velocity | meas | m_MR_Vmax_Vmax | unilateral | 125200\20351-3\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | MR Vmean | Unsigned Time Average Peak Vel | meas | m_MR_VTI_CardTAPV | unilateral | 125200\20352-1\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 20352-1, Time Averaged Mean Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | MR Volume | none | calc | c_MR_Vol | unilateral | 125200\125216\33878-0\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 33878-0, Volume Flow |  | UCUM, ml, Milliliter |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Adult Echocardiography Procedure Report | Adult Echo | MR VTI | Velocity Time Integral | meas | m_MR_VTI_VTI | unilateral | 125200\20354-7\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | MV A Dur Time | Time | meas | m_MVA_Dur_DopTime | unilateral | 125200\G-0385\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-0385, Mitral Valve A-Wave Duration |  | UCUM, sec, Seconds
Adult Echocardiography Procedure Report | Adult Echo | MV A-C Interval Time | Time | meas | m_MV_AC_Int_MMTime | unilateral | 125200\C12207-04\G-0394\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-04, Mitral Valve A-C Interval |  | UCUM, sec, Seconds |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | MV Acc Time | Time | meas | m_MV_AccelTime_DopTime | unilateral | 125200\20168-1\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV Acc Time Slope | Doppler Slope | meas | m_MV_AccelTime_DopSlope | unilateral | 125200\20167-3\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20167-3, Acceleration Slope |  | UCUM, cm/s2, Centimeter Per Second Square |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV Alias Vel | Unsigned Maximum Velocity | meas | m_MV_AliasVel_2DVel | unilateral | 125200\59130-5\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 59130-5, Alias velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV Area (Planim) | Area | meas | m_MV_Area_Area | unilateral | 125200\125220\G-038E\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125220, Planimetry
Adult Echocardiography Procedure Report | Adult Echo | MV D-E Exc Dist | Length | meas | m_MV_DE_Exc_Distance | unilateral | 125200\C12207-01\G-0394\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-01, Mitral Valve D-E Excursion |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | MV D-E Slope | MMode Slope | meas | m_MV_DE_Slope_MMSlope | unilateral | 125200\C12207-02\G-0394\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-02, Mitral Valve D-E Slope |  | UCUM, cm/s, Centimeter Per Second |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | MV Dec Slope | Doppler Slope | meas | m_MV_DecSlope_DopSlope | unilateral | 125200\20216-8\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20216-8, Deceleration Slope |  | UCUM, cm/s2, Centimeter Per Second Square |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV Dec Time | Time | meas | m_MV_DecelTime_DopTime | unilateral | 125200\20217-6\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20217-6, Deceleration Time |  | UCUM, sec, Seconds |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV Diam | Length | meas | m_MV_Diam_Distance | unilateral | 125200\G-038F\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038F, Cardiovascular Orifice Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | MV E/A | none | calc | c_MV_EA | unilateral | 125200\18038-0\T-35300 | SRT, T-35300, Mitral Valve | LN, 18038-0, Mitral Valve E to A Ratio |  | UCUM, 1, no units
Adult Echocardiography Procedure Report | Adult Echo | MV E-E Sep | Length | meas | m_MV_EE_Separ_Distance | unilateral | 125200\59098-4\G-0394\T-35300 | SRT, T-35300, Mitral Valve | LN, 59098-4, Mitral Valve E-septal Separation |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | MV E-F Slope | MMode Slope | meas | m_MV_EF_Slope_MMSlope | unilateral | 125200\18040-6\G-0394\T-35300 | SRT, T-35300, Mitral Valve | LN, 18040-6, Mitral Valve E-F Slope by M-Mode |  | UCUM, cm/s, Centimeter Per Second |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | MV EPSS | Length | meas | m_MV_EPSS_Distance | unilateral | 125200\18036-4\G-0394\T-35300 | SRT, T-35300, Mitral Valve | LN, 18036-4, Mitral Valve EPSS, E wave |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | MV Max PG | Maximum PG | meas | m_MV_Vmax_MaxPG | unilateral | 125200\20247-3\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV Mean PG | Mean PG | meas | m_MV_VTI_CardPGmean | unilateral | 125200\20256-4\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV P1/2t | Pressure Half Time | meas | m_MV_PHalfT_PHalftime | unilateral | 125200\20280-4\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20280-4, Pressure Half-Time |  | UCUM, ms, Millisecond |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV P1/2t Vmax | Unsigned Maximum Velocity | meas | m_MV_PHalfT_Vmax | unilateral | 125200\C12222-03\R-42047\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12222-03, Pressure Half-Time Peak velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV Peak A Vel | Unsigned Maximum Velocity | meas | m_MV_PeakAVel_Vel | unilateral | 125200\17978-8\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 17978-8, Mitral Valve A-Wave Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV Peak E Vel | Unsigned Maximum Velocity | meas | m_MV_PeakEVel_Vel | unilateral | 125200\18037-2\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 18037-2, Mitral Valve E-Wave Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV Radius | Length | meas | m_MV_Radius_Distance | unilateral | 125200\59102-4\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 59102-4, Flow Radius |  | UCUM, cm, Centimeter |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV R-R | Time | meas | m_HR_MV_HeartRate | unilateral | 125200\8867-4\T-35300 | SRT, T-35300, Mitral Valve | LN, 8867-4, Heart Rate |  | UCUM, sec, Seconds
Adult Echocardiography Procedure Report | Adult Echo | MV Vmax | Unsigned Maximum Velocity | meas | m_MV_Vmax_Vmax | unilateral | 125200\20351-3\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV Vmean | Unsigned Time Average Peak Vel | meas | m_MV_VTI_CardTAPV | unilateral | 125200\20352-1\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20352-1, Time Averaged Mean Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MV VTI | Velocity Time Integral | meas | m_MV_VTI_VTI | unilateral | 125200\20354-7\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | MVA (P1/2t) | none | calc | c_MVA_PHalfT | unilateral | 125200\125210\G-038E\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | [5 empty] | DCM, 125210, Area by Pressure Half-Time
Adult Echocardiography Procedure Report | Adult Echo | MVA (PISA) | none | calc | c_MVA_PISA | unilateral | 125200\125216\G-038E\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | [5 empty] | DCM, 125216, Proximal Isovelocity Surface Area
Adult Echocardiography Procedure Report | Adult Echo | MVA (VTI) | none | calc | c_MVA_VTI | unilateral | 125200\125215\G-038E\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | [5 empty] | DCM, 125215, Continuity Equation by Velocity Time Integral
Adult Echocardiography Procedure Report | Adult Echo | Nonplanar Angle | none | meas | m_mvq_NonplnAngle | unilateral | 125200\C12207-20\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-20, Mitral Valve Nonplanar Angle |  | UCUM, deg, deg |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | P1 Prolapse Vol | none | meas | m_mvq_P1ProVol | unilateral | 125200\C12207-37\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-37, Mitral Valve P1 Prolapse Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | P1 Tent Vol | none | meas | m_mvq_P1TentVol | unilateral | 125200\C12207-36\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-36, Mitral Valve P1 Tenting Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | P2 Prolapse Vol | none | meas | m_mvq_P2ProVol | unilateral | 125200\C12207-39\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-39, Mitral Valve P2 Prolapse Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | P2 Tent Vol | none | meas | m_mvq_P2TentVol | unilateral | 125200\C12207-38\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-38, Mitral Valve P2 Tenting Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | P3 Prolapse Vol | none | meas | m_mvq_P3ProVol | unilateral | 125200\C12207-43\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-43, Mitral Valve P3 Prolapse Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | P3 Tent Vol | none | meas | m_mvq_P3TentVol | unilateral | 125200\C12207-40\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-40, Mitral Valve P3 Tenting Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | PI End Dias PG | Maximum PG | meas | m_PI_EndDiVel_PG | unilateral | 125200\20247-3\G-0367\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | PI End Dias Vel | Unsigned Maximum Velocity | meas | m_PI_EndDiVel_Vel | unilateral | 125200\11653-3\G-0367\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | PISA (AI) | none | calc | c_PISA_AI | unilateral | 125200\125216\C12211-01\G-0367\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12211-01, Aortic Valve Flow Area |  | UCUM, cm2, Square Centimeter |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Adult Echocardiography Procedure Report | Adult Echo | PISA (MR) | none | calc | c_PISA_MR | unilateral | 125200\125216\C12207-06\G-0367\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-06, Mitral Valve Flow Area |  | UCUM, cm2, Square Centimeter |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Adult Echocardiography Procedure Report | Adult Echo | PISA (TR) | none | calc | c_PISA_TR | unilateral | 125200\125216\C12208-05\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-05, Tricuspid Valve Flow Area |  | UCUM, cm2, Square Centimeter |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Adult Echocardiography Procedure Report | Adult Echo | Post Leaflet | none | meas | m_mvq_PostLeaf | unilateral | 125200\C12207-11\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-11, Mitral Valve Posterior Leaflet Length |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | Post Leaflet Angle | none | meas | m_mvq_PLeafAngle | unilateral | 125200\C12207-19\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-19, Mitral Valve Posterior Leaflet Angle |  | UCUM, deg, deg |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | Post Leaflets 3D Area | none | meas | m_mvq_PostLeaf3DA | unilateral | 125200\C12207-13\C12228-08\P0-02242\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-13, Mitral Valve Posterior Leaflets 3D Area |  | UCUM, cm2, Square Centimeter |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | Post Med Chord | none | meas | m_mvq_PMChord | unilateral | 125200\C12207-17\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-17, Mitral Valve Posteromedial Chord Length |  | UCUM, mm, Millimeter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | Prolapse Height | none | meas | m_mvq_ProlHeight | unilateral | 125200\C12207-24\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-24, Mitral Valve Prolapse Height |  | UCUM, mm, Millimeter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | Prolapse Vol | none | meas | m_mvq_ProVol | unilateral | 125200\C12207-15\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-15, Mitral Valve Prolapse Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | Pulm A Revs Dur Time | Time | meas | m_PulmARevsDur_DopTime | unilateral | 125200\G-038B\T-48581 | SRT, T-48581, Pulmonary Venous Structure | SRT, G-038B, Pulmonary Vein A-Wave Duration |  | UCUM, sec, Seconds
Adult Echocardiography Procedure Report | Adult Echo | Pulm A Revs Vel | Unsigned Maximum Velocity | meas | m_PulmARevsVel_Vel | unilateral | 125200\29453-8\T-48581 | SRT, T-48581, Pulmonary Venous Structure | LN, 29453-8, Pulmonary Vein Atrial Contraction Reversal Peak Velocity |  | UCUM, cm/s, Centimeter Per Second
Adult Echocardiography Procedure Report | Adult Echo | Pulm Dias Vel | Unsigned Maximum Velocity | meas | m_PulmDiasVel_Vel | unilateral | 125200\29451-2\T-48581 | SRT, T-48581, Pulmonary Venous Structure | LN, 29451-2, Pulmonary Vein Diastolic Peak Velocity |  | UCUM, cm/s, Centimeter Per Second
Adult Echocardiography Procedure Report | Adult Echo | Pulm S/D | none | calc | c_PulmSD | unilateral | 125200\29452-0\T-48581 | SRT, T-48581, Pulmonary Venous Structure | LN, 29452-0, Pulmonary Vein Systolic to Diastolic Ratio |  | UCUM, 1, no units
Adult Echocardiography Procedure Report | Adult Echo | Pulm Sys Vel | Unsigned Maximum Velocity | meas | m_PulmSysVel_Vel | unilateral | 125200\29450-4\T-48581 | SRT, T-48581, Pulmonary Venous Structure | LN, 29450-4, Pulmonary Vein Systolic Peak Velocity |  | UCUM, cm/s, Centimeter Per Second
Adult Echocardiography Procedure Report | Adult Echo | PV Acc Time | Time | meas | m_PV_AccelTime_DopTime | unilateral | 125200\20168-1\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | PV Acc Time Slope | Doppler Slope | meas | m_PV_AccelTime_DopSlope | unilateral | 125200\20167-3\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20167-3, Acceleration Slope |  | UCUM, cm/s2, Centimeter Per Second Square |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | PV Max PG | Maximum PG | meas | m_PV_Vmax_MaxPG | unilateral | 125200\20247-3\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | PV Mean PG | Mean PG | meas | m_PV_VTI_CardPGmean | unilateral | 125200\20256-4\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | PV R-R | Time | meas | m_HR_PV_HeartRate | unilateral | 125200\8867-4\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 8867-4, Heart Rate |  | UCUM, sec, Seconds
Adult Echocardiography Procedure Report | Adult Echo | PV Vmax | Unsigned Maximum Velocity | meas | m_PV_Vmax_Vmax | unilateral | 125200\20351-3\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | PV Vmean | Unsigned Time Average Peak Vel | meas | m_PV_VTI_CardTAPV | unilateral | 125200\20352-1\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20352-1, Time Averaged Mean Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | PV VTI | Velocity Time Integral | meas | m_PV_VTI_VTI | unilateral | 125200\20354-7\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | PVA (Vmax) | none | calc | c_PVA_Vmax | unilateral | 125200\125214\G-038E\T-35200 | SRT, T-35200, Pulmonic Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | [5 empty] | DCM, 125214, Continuity Equation by Peak Velocity
Adult Echocardiography Procedure Report | Adult Echo | PVA (VTI) | none | calc | c_PVA_VTI | unilateral | 125200\125215\G-038E\T-35200 | SRT, T-35200, Pulmonic Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | [5 empty] | DCM, 125215, Continuity Equation by Velocity Time Integral
Adult Echocardiography Procedure Report | Adult Echo | Qp/Qs | none | calc | c_QpQs | unilateral | 125200\29462-9\P5-30031 | SRT, P5-30031, Cardiac Shunt Study | LN, 29462-9, Pulmonary-to-Systemic Shunt Flow Ratio |  | UCUM, 1, no units
Adult Echocardiography Procedure Report | Adult Echo | R to AV Closure | Time | meas | m_RtoAVc_DopTime | unilateral | 125200\C12211-07\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12211-07, R Wave to Aortic Valve Closure Time |  | UCUM, ms, Millisecond
Adult Echocardiography Procedure Report | Adult Echo | R to AV Open | Time | meas | m_RtoAVo_DopTime | unilateral | 125200\C12211-06\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12211-06, R Wave to Aortic Valve Opening Time |  | UCUM, ms, Millisecond
Adult Echocardiography Procedure Report | Adult Echo | R to MV Closure | Time | meas | m_RtoMVc_DopTime | unilateral | 125200\C12207-42\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-42, R Wave to Mitral Valve Closure Time |  | UCUM, ms, Millisecond
Adult Echocardiography Procedure Report | Adult Echo | R to MV Open | Time | meas | m_RtoMVo_DopTime | unilateral | 125200\C12207-41\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-41, R Wave to Mitral Valve Opening Time |  | UCUM, ms, Millisecond
Adult Echocardiography Procedure Report | Adult Echo | RA Pressure | Pressure | meas | m_RA_Press_UserPressure | unilateral | 125200\18070-3\T-32200 | SRT, T-32200, Right Atrium | LN, 18070-3, Right Atrium Systolic Pressure |  | UCUM, mm[Hg], Millimeters Of Mercury
Adult Echocardiography Procedure Report | Adult Echo | RPA Diam | Length | meas | m_RPA_Diam_Distance | unilateral | 125200\18021-6\G-03A2\T-44000 | SRT, T-44000, Pulmonary artery | LN, 18021-6, Right Pulmonary Artery Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | R-R Interval | Time | meas | m_R_R_INT | unilateral | 125200\8867-4\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 8867-4, Heart Rate |  | UCUM, sec, Seconds |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | R-R time | none | meas | m_3DQA_RRTime | unilateral | 125200\C12203-26\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-26, Time between R-R waves |  | UCUM, ms, Millisecond |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | RV EDA (PSAX) | Area | meas | m_RVEDA_2D_Area | unilateral | 125200\C12204-01\F-32011\G-039B\G-03A2\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-01, Right Ventricular Diastolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-039B, Parasternal short axis at the Papillary Muscle level | SRT, F-32011, End Diastole
Adult Echocardiography Procedure Report | Adult Echo | RV ESA (PSAX) | Area | meas | m_RVESA_2D_Area | unilateral | 125200\C12204-02\G-039B\G-03A2\R-FAB5B\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-02, Right Ventricular Systolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-039B, Parasternal short axis at the Papillary Muscle level | SRT, R-FAB5B, End Systole
Adult Echocardiography Procedure Report | Adult Echo | RV ET Time | Time | meas | m_RV_Eject_Tm_MMTime | unilateral | 125200\20222-6\G-0394\T-32500 | SRT, T-32500, Right Ventricle | LN, 20222-6, Ejection Time |  | UCUM, sec, Seconds |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | RV FAC (PSAX) | none | calc | c_RV_FAC | unilateral | 125200\C12204-03\G-039B\G-03A2\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-03, Right Ventricular Fractional Area Change |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-039B, Parasternal short axis at the Papillary Muscle level
Adult Echocardiography Procedure Report | Adult Echo | RV PEP Time | Time | meas | m_RV_PreEject_Tm_MMTime | unilateral | 125200\59085-1\G-0394\T-32500 | SRT, T-32500, Right Ventricle | LN, 59085-1, Pre-Ejection Period |  | UCUM, sec, Seconds |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | RV PEP/ET | none | calc | c_RV_PEP_ET | unilateral | 125200\C12203-04\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12203-04, PEP/ET |  | UCUM, 1, no units
Adult Echocardiography Procedure Report | Adult Echo | RV/LV | none | calc | c_RV_LV_Ratio | unilateral | 125200\C12204-04\G-03A2\T-32500 | SRT, T-32500, Right Ventricle | 99PMSBLUS, C12204-04, Right and Left Ventricular End Diastolic Diameter Ratio |  | UCUM, 1, no units |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | RVAWd (2D) | Length | meas | m_RVAWd_2D_Distance | unilateral | 125200\18153-7\G-03A2\T-32500 | SRT, T-32500, Right Ventricle | LN, 18153-7, Right Ventricular Anterior Wall Diastolic Thickness |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | RVAWd (MM) | Length | meas | m_RVAWd_MM_Distance | unilateral | 125200\18153-7\G-0394\T-32500 | SRT, T-32500, Right Ventricle | LN, 18153-7, Right Ventricular Anterior Wall Diastolic Thickness |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | RVIDd (2D) | Length | meas | m_RVIDd_2D_Distance | unilateral | 125200\20304-2\G-03A2\T-32500 | SRT, T-32500, Right Ventricle | LN, 20304-2, Right Ventricular Internal Diastolic Dimension |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | RVIDd (MM) | Length | meas | m_RVIDd_MM_Distance | unilateral | 125200\20304-2\G-0394\T-32500 | SRT, T-32500, Right Ventricle | LN, 20304-2, Right Ventricular Internal Diastolic Dimension |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | RVOT Area | none | calc | c_RVOT_Area | unilateral | 125200\G-038E\G-03A2\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | SRT, T-32550, Right Ventricle Outflow Tract | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | RVOT Diam | Length | meas | m_RVOT_Diam_Distance | unilateral | 125200\G-038F\G-03A2\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, G-038F, Cardiovascular Orifice Diameter |  | UCUM, cm, Centimeter | SRT, T-32550, Right Ventricle Outflow Tract | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | RVOT Max PG | Maximum PG | meas | m_RVOT_Vmax_MaxPG | unilateral | 125200\20247-3\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury | SRT, T-32550, Right Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | RVOT Mean PG | Mean PG | meas | m_RVOT_VTI_CardPGmean | unilateral | 125200\20256-4\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury | SRT, T-32550, Right Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | RVOT Vmax | Unsigned Maximum Velocity | meas | m_RVOT_Vmax_Vmax | unilateral | 125200\20351-3\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, T-32550, Right Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | RVOT Vmean | Unsigned Time Average Peak Vel | meas | m_RVOT_VTI_CardTAPV | unilateral | 125200\20352-1\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20352-1, Time Averaged Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, T-32550, Right Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | RVOT VTI | Velocity Time Integral | meas | m_RVOT_VTI_VTI | unilateral | 125200\20354-7\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | SRT, T-32550, Right Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | RVSP | none | calc | c_RVSP | unilateral | 125200\G-0380\T-32500 | SRT, T-32500, Right Ventricle | SRT, G-0380, Right Ventricular Peak Systolic Pressure |  | UCUM, mm[Hg], Millimeters Of Mercury
Adult Echocardiography Procedure Report | Adult Echo | SAX A Strain (aCMQ) | none | meas | m_CMQ_Str_SAXA | unilateral | 125200\C12203-66\C12226-05\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-66, Strain peak value of the Left Ventricle |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-05, Short Axis Apical |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | SAX B Strain (aCMQ) | none | meas | m_CMQ_Str_SAXB | unilateral | 125200\C12203-66\C12226-07\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-66, Strain peak value of the Left Ventricle |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-07, Short Axis Basal |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | SAX EDA (a2DQ) | none | meas | m_2DQ_EDA_SAX_S_AR_SM | unilateral | 125200\C12226-06\C12228-09\C99-12228-05\G-0375\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0375, Left Ventricular Diastolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-06, Short Axis Middle |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | SAX EDA (aCMQ) | Area | meas | m_CMQ_SAXSPSM_EDA | unilateral | 125200\C12226-06\C12228-10\C99-12228-07\G-0375\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0375, Left Ventricular Diastolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-06, Short Axis Middle |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | SAX ESA (a2DQ) | none | meas | m_2DQ_ESA_SAX_S_AR_SM | unilateral | 125200\C12226-06\C12228-09\C99-12228-05\G-0374\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0374, Left Ventricular Systolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-06, Short Axis Middle |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | SAX ESA (aCMQ) | Area | meas | m_CMQ_SAXSPSM_ESA | unilateral | 125200\C12226-06\C12228-10\C99-12228-07\G-0374\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0374, Left Ventricular Systolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-06, Short Axis Middle |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | SAX FAC (a2DQ) | none | meas | m_2DQ_FAC_SAX_S_AR_SM | unilateral | 125200\C12226-06\C12228-09\C99-12228-05\G-0376\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0376, Left Ventricular Fractional Area Change |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-06, Short Axis Middle |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-05, QLab aEF FAC Method
Adult Echocardiography Procedure Report | Adult Echo | SAX FAC (aCMQ) | none | meas | m_CMQ_SAXSPSM_FAC | unilateral | 125200\C12226-06\C12228-10\C99-12228-07\G-0376\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0376, Left Ventricular Fractional Area Change |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-06, Short Axis Middle |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | SAX M Strain (aCMQ) | none | meas | m_CMQ_Str_SAXM | unilateral | 125200\C12203-66\C12226-06\C12228-10\C99-12228-07\G-03A2\T-D0772 | SRT, T-D0772, Myocardial Wall | 99PMSBLUS, C12203-66, Strain peak value of the Left Ventricle |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | 99PMSBLUS, C12226-06, Short Axis Middle |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-07, QLab Global Method
Adult Echocardiography Procedure Report | Adult Echo | Score | none | meas | #Score | n/a | 125200\G-C1E3 |  | SRT, G-C1E3, Score
Adult Echocardiography Procedure Report | Adult Echo | SI (2D-Cubed) | none | calc | c_SI_Cubed | unilateral | 125200\125206\F-00078\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | SI (2D-Teich) | none | calc | c_SI_Teich | unilateral | 125200\125209\F-00078\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | SI (A/L) | none | calc | c_SI_SPE | unilateral | 125200\125226\F-00078\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | SI (MM-Cubed) | none | calc | c_SI_Cubed_MM | unilateral | 125200\125206\F-00078\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  | UCUM, ml/m2, ml/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | SI (MM-Teich) | none | calc | c_SI_Teich_MM | unilateral | 125200\125209\F-00078\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  | UCUM, ml/m2, ml/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | SI(MOD-bp) | none | calc | c_SI_MOD_BP | unilateral | 125200\125207\F-00078\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Adult Echocardiography Procedure Report | Adult Echo | SI(MOD-sp2) | none | calc | c_SI_MOD_SP2 | unilateral | 125200\125208\F-00078\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | SI(MOD-sp4) | none | calc | c_SI_MOD_SP4 | unilateral | 125200\125208\F-00078\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | Stroke Vol | none | meas | m_3DQ_SV | unilateral | 125200\C12228-07\F-32120\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | SV (2D-Cubed) | none | calc | c_SV_Cubed | unilateral | 125200\125206\F-32120\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | SV (2D-Teich) | none | calc | c_SV_Teich | unilateral | 125200\125209\F-32120\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | SV (A/L) | none | calc | c_SV_SPE | unilateral | 125200\125226\F-32120\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Adult Echocardiography Procedure Report | Adult Echo | SV (LVOT) | none | calc | c_SV_LVOT | unilateral | 125200\F-32120\T-32600\T-32650 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter | SRT, T-32650, Left Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | SV (MM-Cubed) | none | calc | c_SV_Cubed_MM | unilateral | 125200\125206\F-32120\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Adult Echocardiography Procedure Report | Adult Echo | SV (MM-Teich) | none | calc | c_SV_Teich_MM | unilateral | 125200\125209\F-32120\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Adult Echocardiography Procedure Report | Adult Echo | SV (MV) | none | calc | c_SV_MV | unilateral | 125200\F-32120\T-35300 | SRT, T-35300, Mitral Valve | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter
Adult Echocardiography Procedure Report | Adult Echo | SV (RVOT) | none | calc | c_SV_RVOT | unilateral | 125200\F-32120\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter | SRT, T-32550, Right Ventricle Outflow Tract
Adult Echocardiography Procedure Report | Adult Echo | SV (TV) | none | calc | c_SV_TV | unilateral | 125200\F-32120\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter
Adult Echocardiography Procedure Report | Adult Echo | SV(MOD-bp) | none | calc | c_SV_MOD_BP | unilateral | 125200\125207\F-32120\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Adult Echocardiography Procedure Report | Adult Echo | SV(MOD-sp2) | none | calc | c_SV_MOD_SP2 | unilateral | 125200\125208\F-32120\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | SV(MOD-sp4) | none | calc | c_SV_MOD_SP4 | unilateral | 125200\125208\F-32120\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Adult Echocardiography Procedure Report | Adult Echo | Tei Index | none | calc | c_TeiIndex | unilateral | 125200\59099-2\T-35300 | SRT, T-35300, Mitral Valve | LN, 59099-2, Myocardial Performance Index (Tei) |  | UCUM, 1, no units
Adult Echocardiography Procedure Report | Adult Echo | Tenting Height | none | meas | m_mvq_OriArea | unilateral | 125200\C12207-29\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-29, Mitral Valve Regurgitant Orifice Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | Tenting Height | none | meas | m_mvq_TentHeight | unilateral | 125200\C12207-23\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-23, Mitral Valve Tenting Height |  | UCUM, mm, Millimeter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | Tenting Vol | none | meas | m_mvq_TentVol | unilateral | 125200\C12207-14\C12228-08\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-14, Mitral Valve Tenting Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | 99PMSBLUS, C12228-08, QLab MVQ
Adult Echocardiography Procedure Report | Adult Echo | Time to Lat E` | Time | meas | m_TimeLatPeakE | unilateral | 125200\59096-8\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59096-8, Time to Left Ventricle E Tissue Velocity |  | UCUM, sec, Seconds | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Time to Lat S | Time | meas | m_TimeLatPeakS | unilateral | 125200\59095-0\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59095-0, Time to Left Ventricle S Tissue Velocity |  | UCUM, sec, Seconds | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Time to Med E` | Time | meas | m_TimeMedPeakE | unilateral | 125200\59096-8\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59096-8, Time to Left Ventricle E Tissue Velocity |  | UCUM, sec, Seconds | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | Time to Med S | Time | meas | m_TimeMedPeakS | unilateral | 125200\59095-0\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59095-0, Time to Left Ventricle S Tissue Velocity |  | UCUM, sec, Seconds | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Adult Echocardiography Procedure Report | Adult Echo | TMAD AP2 M MV (a2DQ) | none | meas | m_2DQ_cakmid_ap2_mv | unilateral | 125200\C12203-62\C12228-09\C99-12228-06\G-03A2\G-A19B\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12203-62, Displacement Midpoint of L/R Mitral Annular Hingepoints to Apex |  | UCUM, mm, Millimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TMAD AP2 M MV (a2DQ) | none | meas | m_2DQ_cakmidnm_ap2_mv | unilateral | 125200\C12203-63\C12228-09\C99-12228-06\G-03A2\G-A19B\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12203-63, Normalized Midpoint of L/R MV Annular Hingepoints to Apex |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TMAD AP2 M MV (aCMQ) | Length | meas | m_CMQ_cakmid_ap2_mv | unilateral | 125200\C12203-62\C99-12228-02\G-03A2\G-A19B\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12203-62, Displacement Midpoint of L/R Mitral Annular Hingepoints to Apex |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | [8 empty] | 99PMSBLUS, C99-12228-02, QLab Simple Method
Adult Echocardiography Procedure Report | Adult Echo | TMAD AP2 M MV (aCMQ) | none | meas | m_CMQ_cakmidnm_ap2_mv | unilateral | 125200\C12203-63\C12228-10\C99-12228-06\G-03A2\G-A19B\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12203-63, Normalized Midpoint of L/R MV Annular Hingepoints to Apex |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TMAD AP2 MV 1 (a2DQ) | none | meas | m_2DQ_cak_ap2_mv1 | unilateral | 125200\C12203-65\C12228-09\C99-12228-06\G-03A2\G-A19B\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12203-65, Displacement of the Left Mitral Annular Hingepoint to the Apex |  | UCUM, mm, Millimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TMAD AP2 MV 1 (aCMQ) | Length | meas | m_CMQ_cak_ap2_mv1 | unilateral | 125200\C12203-65\C12228-10\C99-12228-06\G-03A2\G-A19B\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12203-65, Displacement of the Left Mitral Annular Hingepoint to the Apex |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TMAD AP2 MV 2 (a2DQ) | none | meas | m_2DQ_cak_ap2_mv2 | unilateral | 125200\C12203-64\C12228-09\C99-12228-06\G-03A2\G-A19B\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12203-64, Displacement of the Right Mitral Annular Hingepoint to Apex |  | UCUM, mm, Millimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TMAD AP4 M MV (a2DQ) | none | meas | m_2DQ_cakmid_ap4_mv | unilateral | 125200\C12203-62\C12228-09\C99-12228-06\G-03A2\G-A19C\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12203-62, Displacement Midpoint of L/R Mitral Annular Hingepoints to Apex |  | UCUM, mm, Millimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TMAD AP4 M MV (a2DQ) | none | meas | m_2DQ_cakmidnm_ap4_mv | unilateral | 125200\C12203-63\C12228-09\C99-12228-06\G-03A2\G-A19C\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12203-63, Normalized Midpoint of L/R MV Annular Hingepoints to Apex |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TMAD AP4 M MV (aCMQ) | Length | meas | m_CMQ_cakmid_ap4_mv | unilateral | 125200\C12203-62\C12228-10\C99-12228-06\G-03A2\G-A19C\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12203-62, Displacement Midpoint of L/R Mitral Annular Hingepoints to Apex |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TMAD AP4 M MV (aCMQ) | none | meas | m_CMQ_cakmidnm_ap4_mv | unilateral | 125200\C12203-63\C12228-10\C99-12228-06\G-03A2\G-A19C\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12203-63, Normalized Midpoint of L/R MV Annular Hingepoints to Apex |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TMAD AP4 MV 1 (a2DQ) | none | meas | m_2DQ_cak_ap4_mv1 | unilateral | 125200\C12203-65\C12228-09\C99-12228-06\G-03A2\G-A19C\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12203-65, Displacement of the Left Mitral Annular Hingepoint to the Apex |  | UCUM, mm, Millimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TMAD AP4 MV 1 (aCMQ) | Length | meas | m_CMQ_cak_ap4_mv1 | unilateral | 125200\C12203-65\C12228-10\C99-12228-06\G-03A2\G-A19C\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12203-65, Displacement of the Left Mitral Annular Hingepoint to the Apex |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TMAD AP4 MV 2 (a2DQ) | none | meas | m_2DQ_cak_ap4_mv2 | unilateral | 125200\C12203-64\C12228-09\C99-12228-06\G-03A2\G-A19C\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12203-64, Displacement of the Right Mitral Annular Hingepoint to Apex |  | UCUM, mm, Millimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-09, QLab a2DQ | [5 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | TMAD AP4 MV 2 (aCMQ) | Length | meas | m_CMQ_cak_ap4_mv2 | unilateral | 125200\C12203-64\C12228-10\C99-12228-06\G-03A2\G-A19C\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12203-64, Displacement of the Right Mitral Annular Hingepoint to Apex |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | 99PMSBLUS, C12228-10, QLab aCMQ | [5 empty] | 99PMSBLUS, C99-12228-06, QLab aTMAD Method
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 12-DIF (%) | none | meas | m_nTmsv_12_MxDiff | unilateral | 125200\C12203-22\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-22, Nrm. Mx Dif of Time to min sys vol for 6 basal N 6 Mid segments |  | UCUM, %, Percent |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 12-DIF (ms) | none | meas | m_Tmsv_12_MxDiff | unilateral | 125200\C12203-14\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-14, Max Dif of Time to min sys vol for 6 basal N 6 Mid segments |  | UCUM, ms, Millisecond |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 12-SD (%) | none | meas | m_nTmsv_12_SD | unilateral | 125200\C12203-19\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-19, Nrm. SD of Time to min sys vol for 6 basal N 6 Mid segments |  | UCUM, %, Percent |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 12-SD (ms) | none | meas | m_Tmsv_12_SD | unilateral | 125200\C12203-11\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-11, SD of Time to min sys vol for 6 basal N 6 Mid segments |  | UCUM, ms, Millisecond |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 16-DIF (%) | none | meas | m_nTmsv_16_MxDiff | unilateral | 125200\C12203-21\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-21, Nrm. Mx Dif of Time to min sys vol for 16 wall segments |  | UCUM, %, Percent |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 16-DIF (ms) | none | meas | m_Tmsv_16_MxDiff | unilateral | 125200\C12203-13\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-13, Max Dif of Time to min sys vol for 16 wall segments |  | UCUM, ms, Millisecond |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 16-SD (%) | none | meas | m_nTmsv_16_SD | unilateral | 125200\C12203-18\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-18, Nrm. SD of Time to min sys vol for 16 wall segments |  | UCUM, %, Percent |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 16-SD (ms) | none | meas | m_Tmsv_16_SD | unilateral | 125200\C12203-10\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-10, SD of Time to min sys vol for 16 wall segments |  | UCUM, ms, Millisecond |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 6-DIF (%) | none | meas | m_nTmsv_6_MxDiff | unilateral | 125200\C12203-23\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-23, Nrm. Mx Dif of Time to min sys vol for 6 basal segments |  | UCUM, %, Percent |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 6-DIF (ms) | none | meas | m_Tmsv_6_MxDiff | unilateral | 125200\C12203-15\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-15, Max Dif of Time to min sys vol for 6 basal segments |  | UCUM, ms, Millisecond |  | SRT, P0-02242, 3D mode
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 6-SD (%) | none | meas | m_nTmsv_6_SD | unilateral | 125200\C12203-20\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-20, Nrm. SD of Time to min sys vol for 6 basal segments |  | UCUM, %, Percent |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv 6-SD (ms) | none | meas | m_Tmsv_6_SD | unilateral | 125200\C12203-12\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-12, SD of Time to min sys vol for 6 basal segments |  | UCUM, ms, Millisecond |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv S-L (%) | none | meas | m_nTmsv_SL_MxDiff | unilateral | 125200\C12203-25\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-25, Nrm. Diff of Time to min sys vol btn basal sept N lat segments |  | UCUM, %, Percent |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv S-L (ms) | none | meas | m_Tmsv_SL_MxDiff | unilateral | 125200\C12203-17\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-17, Dif of Time to min sys vol btn basal sept N lat segments |  | UCUM, ms, Millisecond |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv S-P (%) | none | meas | m_nTmsv_SP_MxDiff | unilateral | 125200\C12203-24\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-24, Nrm. Diff of Time to min sys vol btn basal sept N post segments |  | UCUM, %, Percent |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | Tmsv S-P (ms) | none | meas | m_Tmsv_SP_MxDiff | unilateral | 125200\C12203-16\C12228-07\P0-02242\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-16, Dif of Time to min sys vol btn basal sept N post segments |  | UCUM, ms, Millisecond |  | SRT, P0-02242, 3D mode |  |  |  | 99PMSBLUS, C12228-07, QLab 3DQA
Adult Echocardiography Procedure Report | Adult Echo | TR Alias Vel | Unsigned Maximum Velocity | meas | m_TR_AliasVel_2DVel | unilateral | 125200\59130-5\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59130-5, Alias velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | TR ERO | none | calc | c_TR_ERO | unilateral | 125200\125216\G-0367\G-038E\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Adult Echocardiography Procedure Report | Adult Echo | TR Flow Rate | none | calc | c_TR_FlowRate | unilateral | 125200\34141-2\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 34141-2, Peak Instantaneous Flow Rate |  | UCUM, ml/sec, ml/sec |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | TR Fraction | none | calc | c_TR_Frac | unilateral | 125200\G-0367\G-0390\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-0390, Regurgitant Fraction |  | UCUM, %, Percent |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | TR Max PG | Maximum PG | meas | m_TR_Vmax_MaxPG | unilateral | 125200\20247-3\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | TR Mean PG | Mean PG | meas | m_TR_VTI_CardPGmean | unilateral | 125200\20256-4\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | TR Radius | Length | meas | m_TR_Radius_Distance | unilateral | 125200\59102-4\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59102-4, Flow Radius |  | UCUM, cm, Centimeter |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | TR Vmax | Unsigned Maximum Velocity | meas | m_TR_Vmax_Vmax | unilateral | 125200\20351-3\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | TR Vmean | Unsigned Time Average Peak Vel | meas | m_TR_VTI_CardTAPV | unilateral | 125200\20352-1\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20352-1, Time Averaged Mean Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | TR Volume | none | calc | c_TR_Vol | unilateral | 125200\33878-0\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 33878-0, Volume Flow |  | UCUM, ml, Milliliter |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | TR VTI | Velocity Time Integral | meas | m_TR_VTI_VTI | unilateral | 125200\20354-7\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter |  |  |  |  | SRT, G-0367, Regurgitant Flow
Adult Echocardiography Procedure Report | Adult Echo | TV A-C Interval Time | Time | meas | m_TV_AC_Int_MMTime | unilateral | 125200\C12208-04\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-04, Tricuspid Valve A-C Interval |  | UCUM, sec, Seconds |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | TV Acc Time | Time | meas | m_TV_AccelTime_DopTime | unilateral | 125200\20168-1\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TV Acc Time Slope | Doppler Slope | meas | m_TV_AccelTime_DopSlope | unilateral | 125200\20167-3\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20167-3, Acceleration Slope |  | UCUM, cm/s2, Centimeter Per Second Square |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TV Alias Vel | Unsigned Maximum Velocity | meas | m_TV_AliasVel_2DVel | unilateral | 125200\59130-5\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59130-5, Alias velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TV Area | none | calc | c_TV_Area | unilateral | 125200\G-038E\G-03A2\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | TV D-E Exc Dist | Length | meas | m_TV_DE_Exc_Distance | unilateral | 125200\C12208-01\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-01, Tricuspid Valve D-E Excursion |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | TV D-E Slope | MMode Slope | meas | m_TV_DE_Slope_MMSlope | unilateral | 125200\C12208-02\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-02, Tricuspid Valve D-E Slope |  | UCUM, cm/s, Centimeter Per Second |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | TV Diam | Length | meas | m_TV_Diam_Distance | unilateral | 125200\G-038F\G-03A2\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-038F, Cardiovascular Orifice Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Adult Echocardiography Procedure Report | Adult Echo | TV E/A | none | calc | c_TV_EA | unilateral | 125200\18039-8\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 18039-8, Tricuspid Valve E to A Ratio |  | UCUM, 1, no units
Adult Echocardiography Procedure Report | Adult Echo | TV E-F Slope | MMode Slope | meas | m_TV_EF_Slope_MMSlope | unilateral | 125200\C12208-03\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12208-03, Tricuspid Valve E-F Slope |  | UCUM, cm/s, Centimeter Per Second |  | SRT, G-0394, M mode
Adult Echocardiography Procedure Report | Adult Echo | TV Max PG | Maximum PG | meas | m_TV_Vmax_MaxPG | unilateral | 125200\20247-3\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TV Mean PG | Mean PG | meas | m_TV_VTI_CardPGmean | unilateral | 125200\20256-4\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TV Peak A Vel | Unsigned Maximum Velocity | meas | m_TV_PeakAVel_Vel | unilateral | 125200\18030-7\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 18030-7, Tricuspid Valve A Wave Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TV Peak E Vel | Unsigned Maximum Velocity | meas | m_TV_PeakEVel_Vel | unilateral | 125200\18031-5\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 18031-5, Tricuspid Valve E Wave Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TV Radius | Length | meas | m_TV_Radius_Distance | unilateral | 125200\59102-4\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59102-4, Flow Radius |  | UCUM, cm, Centimeter |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TV R-R | Time | meas | m_HR_TV_HeartRate | unilateral | 125200\8867-4\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 8867-4, Heart Rate |  | UCUM, sec, Seconds
Adult Echocardiography Procedure Report | Adult Echo | TV Vmax | Unsigned Maximum Velocity | meas | m_TV_Vmax_Vmax | unilateral | 125200\20351-3\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TV Vmean | Unsigned Time Average Peak Vel | meas | m_TV_VTI_CardTAPV | unilateral | 125200\20352-1\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20352-1, Time Averaged Mean Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TV VTI | Velocity Time Integral | meas | m_TV_VTI_VTI | unilateral | 125200\20354-7\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter |  |  |  |  | SRT, R-42047, Antegrade Flow
Adult Echocardiography Procedure Report | Adult Echo | TVA (PISA) | none | calc | c_TVA_PISA | unilateral | 125200\125216\G-038E\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | [5 empty] | DCM, 125216, Proximal Isovelocity Surface Area
Adult Echocardiography Procedure Report | Adult Echo | Wall Segment | none | meas | #WallSegment | n/a | 125200\18179-2 |  | LN, 18179-2, Wall Segment
OB-GYN Ultrasound Procedure Report | OB | 2 Beat Pk-to-Pk | Time | meas | MEAS_2_BEAT_PK_TO_PK | unilateral | 125000\ C12019-01\11635-5 | LN, 11635-5, Fetal Heart Reactivity | 99PMSBLUS, C12019-01, Peak-to-Peak time interval over two beats |  | UCUM, sec, Seconds
OB-GYN Ultrasound Procedure Report | OB | AC | Circumference | meas | MEAS_AC_TRACED | unilateral | 125000\11979-2 |  | LN, 11979-2, Abdominal Circumference |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | AC(c) | none | calc | CALC_AC | unilateral | 125000\11979-2 |  | LN, 11979-2, Abdominal Circumference |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | AFI | none | calc | CALC_AFI | unilateral | 125000\11627-7 |  | LN, 11627-7, Amniotic Fluid Index |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Amniotic Fluid Volume | none | meas | #fet_afv | n/a | 125000\11630-1 |  | LN, 11630-1, Amniotic Fluid Volume
OB-GYN Ultrasound Procedure Report | OB | APAD | Length | meas | MEAS_AD_AP | unilateral | 125000\11818-2 |  | LN, 11818-2, Anterior-Posterior Abdominal Diameter |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | APTD | Length | meas | MEAS_APTD | unilateral | 125000\11819-0 |  | LN, 11819-0, Anterior-Posterior Trunk Diameter |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | APTD*TTD | none | calc | CALC_APTDxTTD | unilateral | 125000\C12005-10 |  | 99PMSBLUS, C12005-10, APTD*TTD |  | UCUM, cm2, Square Centimeter
OB-GYN Ultrasound Procedure Report | OB | APThD | Length | meas | MEAS_TD_AP | unilateral | 125000\C12005-07 |  | 99PMSBLUS, C12005-07, Anterior-Posterior Thoracic Diameter |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | AUA | none | calc | CALC_AUA | unilateral | 125000\11888-5 |  | LN, 11888-5, Composite Ultrasound Age
OB-GYN Ultrasound Procedure Report | OB | Biophysical Profile Sum Score | none | meas | #fet_score | n/a | 125000\11634-3 |  | LN, 11634-3, Biophysical Profile Sum Score
OB-GYN Ultrasound Procedure Report | OB | Bladder Height | Length | meas | MEAS_GYN_BLADDER_HEIGHT | unilateral | 125000\C12011-03 |  | 99PMSBLUS, C12011-03, Bladder Height |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Bladder Length | Length | meas | MEAS_GYN_BLADDER_LENGTH | unilateral | 125000\C12011-01 |  | 99PMSBLUS, C12011-01, Bladder Length |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Bladder Vol | none | calc | CALC_GYN_BLADDER_VOL | unilateral | 125000\C12011-04 |  | 99PMSBLUS, C12011-04, Bladder Volume |  | UCUM, cm3, Cubic Centimeter
OB-GYN Ultrasound Procedure Report | OB | Bladder Width | Length | meas | MEAS_GYN_BLADDER_WIDTH | unilateral | 125000\C12011-02 |  | 99PMSBLUS, C12011-02, Bladder Width |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | BPD | Length | meas | MEAS_BPD | unilateral | 125000\11820-8 |  | LN, 11820-8, Biparietal Diameter |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Cerebellum | Length | meas | MEAS_CD | unilateral | 125000\11863-8 |  | LN, 11863-8, Trans Cerebellar Diameter |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Cervix | Contour | meas | MEAS_CERVIX_LENGTH | unilateral | 125000\11961-0 |  | LN, 11961-0, Cervix Length |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | CI(BPD,OFD) | none | calc | CALC_CI_BPD_OFD | unilateral | 125000\11823-2 |  | LN, 11823-2, Cephalic Index |  | UCUM, 1, no units
OB-GYN Ultrasound Procedure Report | OB | Cist Mag | Length | meas | MEAS_CIST_MAG | unilateral | 125000\11860-4 |  | LN, 11860-4, Cisterna Magna length |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Clavicle | Length | meas | MEAS_CLAV | unilateral | 125000\11962-8 |  | LN, 11962-8, Clavicle length |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | CRL | Length | meas | MEAS_CRL | unilateral | 125000\11957-8 |  | LN, 11957-8, Crown Rump Length |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Ductus Venosus AI | Doppler Slope | meas | m_Ductus_Venosus_AS | unilateral | 125000\20167-3\C12141-01 | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square
OB-GYN Ultrasound Procedure Report | OB | Ductus Venosus AT | Time | meas | m_Ductus_Venosus_AT | unilateral | 125000\20168-1\C12141-01 | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds
OB-GYN Ultrasound Procedure Report | OB | Ductus Venosus D/S | none | calc | c_Ductus_Venosus_DS | unilateral | 125000\C12121-01\C12141-01 | 99PMSBLUS, C12141-01, Ductus Venosus | 99PMSBLUS, C12121-01, Diastolic to Systolic Velocity Ratio |  | UCUM, 1, ratio
OB-GYN Ultrasound Procedure Report | OB | Ductus Venosus EDV | Signed End Diastolic Vel | meas | m_Ductus_Venosus_EDV | unilateral | 125000\11653-3\C12141-01 | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second
OB-GYN Ultrasound Procedure Report | OB | Ductus Venosus MDV | Signed Minimum Velocity | meas | m_Ductus_Venosus_MDV | unilateral | 125000\11665-7\C12141-01 | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second
OB-GYN Ultrasound Procedure Report | OB | Ductus Venosus PI | none | calc | c_Ductus_Venosus_PI | unilateral | 125000\12008-9\C12141-01 | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio
OB-GYN Ultrasound Procedure Report | OB | Ductus Venosus PSV | Signed Maximum Velocity | meas | m_Ductus_Venosus_PSV | unilateral | 125000\11726-7\C12141-01 | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second
OB-GYN Ultrasound Procedure Report | OB | Ductus Venosus RI | none | calc | c_Ductus_Venosus_RI | unilateral | 125000\12023-8\C12141-01 | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio
OB-GYN Ultrasound Procedure Report | OB | Ductus Venosus S/D | none | calc | c_Ductus_Venosus_SD | unilateral | 125000\12144-2\C12141-01 | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio
OB-GYN Ultrasound Procedure Report | OB | Ductus Venosus TAMV | Time Average Mean Velocity | meas | m_Ductus_Venosus_TAMV | unilateral | 125000\20352-1\C12141-01 | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 20352-1, Time Averaged Mean Velocity |  | UCUM, cm/s, Centimeter Per Second
OB-GYN Ultrasound Procedure Report | OB | Ductus Venosus TAPV | Signed Time Average Peak Vel | meas | m_Ductus_Venosus_TAPV | unilateral | 125000\11692-1\C12141-01 | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 11692-1, Time Averaged Peak Velocity |  | UCUM, cm/s, Centimeter Per Second
OB-GYN Ultrasound Procedure Report | OB | Ductus Venosus VTI | Velocity Time Integral | meas | m_Ductus_Venosus_VTI | unilateral | 125000\20354-7\C12141-01 | 99PMSBLUS, C12141-01, Ductus Venosus | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Ear | Length | meas | MEAS_EAR | unilateral | 125000\C12005-01 |  | 99PMSBLUS, C12005-01, Ear Length |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | EDV | Signed End Diastolic Vel | meas | m_Lt_UterArt_EDV | unilateral | 125000\11653-3\G-A101\T-46820 | SRT, T-46820, Uterine Artery | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [8 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | EDV | Signed End Diastolic Vel | meas | m_Rt_UterArt_EDV | unilateral | 125000\11653-3\G-A100\T-46820 | SRT, T-46820, Uterine Artery | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [8 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | EFW(AC)Campbell | none | calc | CALC_EFW_AC_CAMPBELL | unilateral | 125000\11756-4 |  | LN, 11756-4, EFW by AC, Campbell 1975 |  | UCUM, g, Gram
OB-GYN Ultrasound Procedure Report | OB | EFW(AC,BPD)Hadl | none | calc | CALC_EFW_AC_BPD_HADLOCK | unilateral | 125000\11738-2 |  | LN, 11738-2, EFW by AC, BPD, Hadlock 1984 |  | UCUM, g, Gram
OB-GYN Ultrasound Procedure Report | OB | EFW(AC,BPD)Sh | none | calc | CALC_EFW_AC_BPD_SHEPHARD | unilateral | 125000\11739-0 |  | LN, 11739-0, EFW by AC and BPD, Shepard 1982 |  | UCUM, g, Gram
OB-GYN Ultrasound Procedure Report | OB | EFW(AC,BPD,FL)Hadl | none | calc | CALC_EFW_AC_BPD_FL_HADLOCK | unilateral | 125000\11735-8 |  | LN, 11735-8, EFW by AC, BPD, FL, Hadlock 1985 |  | UCUM, g, Gram
OB-GYN Ultrasound Procedure Report | OB | EFW(AC,BPD,FL)Shino | none | calc | CALC_EFW_AC_BPD_FL_SHINO | unilateral | 125000\C12014-01 |  | 99PMSBLUS, C12014-01, EFW by AC, BPD and FL,Shinozuka 2000 |  | UCUM, g, Gram
OB-GYN Ultrasound Procedure Report | OB | EFW(AC,FL)Hadl | none | calc | CALC_EFW_AC_FL_HADLOCK | unilateral | 125000\11751-5 |  | LN, 11751-5, EFW by AC, FL, Hadlock 1985 |  | UCUM, g, Gram
OB-GYN Ultrasound Procedure Report | OB | EFW(AC,HC,FL)Hadl | none | calc | CALC_EFW_AC_HC_FL_HADLOCK | unilateral | 125000\11746-5 |  | LN, 11746-5, EFW by AC, FL, HC, Hadlock 1985 |  | UCUM, g, Gram
OB-GYN Ultrasound Procedure Report | OB | EFW(B,APxTTD,F)Shino | none | calc | CALC_EFW_BPD_APTDTTD_FL_SHINO | unilateral | 125000\33143-9 |  | LN, 33143-9, EFW3 by Shinozuka 1996 |  | UCUM, g, Gram
OB-GYN Ultrasound Procedure Report | OB | EFW(B,H,A,F)Hadl | none | calc | CALC_EFW_B_H_A_F_HADLOCK | unilateral | 125000\11732-5 |  | LN, 11732-5, EFW by AC, BPD, FL, HC, Hadlock 1985 |  | UCUM, g, Gram
OB-GYN Ultrasound Procedure Report | OB | EFW(BPD,AD,FL)Tokyo | none | calc | CALC_EFW_BPD_AD_FL_TOKYO | unilateral | 125000\C12014-02 |  | 99PMSBLUS, C12014-02, EFW by BPD, APAD, TAD, FL, Tokyo 1986 |  | UCUM, g, Gram
OB-GYN Ultrasound Procedure Report | OB | EFW(BPD,FTA,FL)Osaka | none | calc | CALC_EFW_BPD_FTA_FL_OSAKA | unilateral | 125000\33140-5 |  | LN, 33140-5, EFW by BPD, FTA, FL, Osaka 1990 |  | UCUM, g, Gram
OB-GYN Ultrasound Procedure Report | OB | EMM | Length | meas | MEAS_EMM | unilateral | 125000\C12009-05 |  | 99PMSBLUS, C12009-05, Endomyometrial Mantle |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Endometrium | Length | meas | MEAS_ENDOMETRIAL_THICKNESS | unilateral | 125000\12145-9 |  | LN, 12145-9, Endometrium Thickness |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Fetal Breathing | none | meas | #fet_breath | n/a | 125000\11632-7 |  | LN, 11632-7, Fetal Breathing
OB-GYN Ultrasound Procedure Report | OB | Fetal Heart Reactivity | none | meas | #fet_heart | n/a | 125000\11635-5 |  | LN, 11635-5, Fetal Heart Reactivity
OB-GYN Ultrasound Procedure Report | OB | Fetal HR | none | calc | CALC_FETAL_HR | unilateral | 125000\11948-7 |  | LN, 11948-7, Fetal Heart Rate |  | UCUM, {H.B.}/min, Beats Per Minute
OB-GYN Ultrasound Procedure Report | OB | Fetal Tone | none | meas | #fet_tone | n/a | 125000\11635-0 |  | LN, 11635-0, Fetal Tone
OB-GYN Ultrasound Procedure Report | OB | Fibula | Length | meas | MEAS_FIB | unilateral | 125000\11964-4 |  | LN, 11964-4, Fibula length |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | FL | Length | meas | MEAS_FL | unilateral | 125000\11963-6 |  | LN, 11963-6, Femur Length |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | FL/AC | none | calc | CALC_FL_AC | unilateral | 125000\11871-1 |  | LN, 11871-1, FL/AC |  | UCUM, 1, ratio
OB-GYN Ultrasound Procedure Report | OB | FL/BPD | none | calc | CALC_FL_BPD | unilateral | 125000\11872-9 |  | LN, 11872-9, FL/BPD |  | UCUM, 1, ratio
OB-GYN Ultrasound Procedure Report | OB | Foot | Length | meas | MEAS_FOOT | unilateral | 125000\11965-1 |  | LN, 11965-1, Foot length |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Frontomaxillary Facial Angle | Alpha Angle | meas | m_frontomax_facial_angle | unilateral | 125000\C12005-11 |  | 99PMSBLUS, C12005-11, Frontomaxillary Facial Angle |  | UCUM, deg, deg
OB-GYN Ultrasound Procedure Report | OB | FTA | Area | meas | MEAS_FTA_TRACED | unilateral | 125000\C12005-02 |  | 99PMSBLUS, C12005-02, Fetal Trunk Cross Sectional Area |  | UCUM, cm2, Square Centimeter
OB-GYN Ultrasound Procedure Report | OB | GA(AC)ASUM(2001) | none | calc | CALC_GA_AC_ASUM2001 | unilateral | 125000\C12013-01 |  | 99PMSBLUS, C12013-01, AC, ASUM 2001
OB-GYN Ultrasound Procedure Report | OB | GA(AC)Hadlock | none | calc | CALC_GA_AC_HADLOCK | unilateral | 125000\11892-7 |  | LN, 11892-7, AC, Hadlock 1984
OB-GYN Ultrasound Procedure Report | OB | GA(AC)Hansmann | none | calc | CALC_GA_AC_HANSMANN | unilateral | 125000\C12013-30 |  | 99PMSBLUS, C12013-30, AC, Hansmann 1986
OB-GYN Ultrasound Procedure Report | OB | GA(AC)Merz(1991) | none | calc | CALC_GA_AC_MERZ | unilateral | 125000\C12013-16 |  | 99PMSBLUS, C12013-16, AC Merz 1991
OB-GYN Ultrasound Procedure Report | OB | GA(AC)Shinozuka | none | calc | CALC_GA_AC_SHINOZUKA | unilateral | 125000\33076-1 |  | LN, 33076-1, AC, Shinozuka 1996
OB-GYN Ultrasound Procedure Report | OB | GA(AC)Verburg | none | calc | CALC_GA_AC_VERBURG | unilateral | 125000\C12013-43 |  | 99PMSBLUS, C12013-43, AC, Verburg 2007
OB-GYN Ultrasound Procedure Report | OB | GA(APTDxTTD)Shino | none | calc | CALC_GA_APTDXTTD_SHINO | unilateral | 125000\33078-7 |  | LN, 33078-7, AxT, Shinozuka 1996
OB-GYN Ultrasound Procedure Report | OB | GA(BPD)ASUM(2001) | none | calc | CALC_GA_BPD_ASUM2001 | unilateral | 125000\C12013-02 |  | 99PMSBLUS, C12013-02, BPD, ASUM 2001
OB-GYN Ultrasound Procedure Report | OB | GA(BPD)Chitty(out-in) | none | calc | CALC_GA_BPD_CHITTY_O_I | unilateral | 125000\33086-0 |  | LN, 33086-0, BPD-oi, Chitty 1997
OB-GYN Ultrasound Procedure Report | OB | GA(BPD)Chitty(out-out) | none | calc | CALC_GA_BPD_CHITTY_O_O | unilateral | 125000\33087-8 |  | LN, 33087-8, BPD-oo, Chitty 1997
OB-GYN Ultrasound Procedure Report | OB | GA(BPD)Hadlock | none | calc | CALC_GA_BPD_HADLOCK | unilateral | 125000\11902-4 |  | LN, 11902-4, BPD, Hadlock 1984
OB-GYN Ultrasound Procedure Report | OB | GA(BPD)Hansmann | none | calc | CALC_GA_BPD_HANSMANN | unilateral | 125000\33538-0 |  | LN, 33538-0, BPD, Hansmann 1986
OB-GYN Ultrasound Procedure Report | OB | GA(BPD)Jeanty | none | calc | CALC_GA_BPD_JEANTY | unilateral | 125000\11905-7 |  | LN, 11905-7, BPD, Jeanty 1984
OB-GYN Ultrasound Procedure Report | OB | GA(BPD)Merz(1991) | none | calc | CALC_GA_BPD_MERZ | unilateral | 125000\C12013-17 |  | 99PMSBLUS, C12013-17, BPD Merz 1991
OB-GYN Ultrasound Procedure Report | OB | GA(BPD)Osaka | none | calc | CALC_GA_BPD_OSAKA | unilateral | 125000\C12013-31 |  | 99PMSBLUS, C12013-31, BPD, Osaka 1988
OB-GYN Ultrasound Procedure Report | OB | GA(BPD)Shinozuka | none | calc | CALC_GA_BPD_SHINO | unilateral | 125000\33084-5 |  | LN, 33084-5, BPD, Shinozuka 1996
OB-GYN Ultrasound Procedure Report | OB | GA(BPD)Tokyo | none | calc | CALC_GA_BPD_TOKYO | unilateral | 125000\C12013-26 |  | 99PMSBLUS, C12013-26, BPD, Tokyo 1989
OB-GYN Ultrasound Procedure Report | OB | GA(BPD)Verburg | none | calc | CALC_GA_BPD_VERBURG | unilateral | 125000\C12013-37 |  | 99PMSBLUS, C12013-37, BPD, Verburg 2008
OB-GYN Ultrasound Procedure Report | OB | GA(Cerebellum)Chitty | none | calc | CALC_GA_CEREBELLUM_CHITTY | unilateral | 125000\C12013-15 |  | 99PMSBLUS, C12013-15, TCD, Chitty 1997
OB-GYN Ultrasound Procedure Report | OB | GA(Cerebellum)Hill | none | calc | CALC_GA_CEREBELLUM_HILL | unilateral | 125000\33134-8 |  | LN, 33134-8, TCD, Hill 1990
OB-GYN Ultrasound Procedure Report | OB | GA(Clavicle)Yarkoni | none | calc | CALC_GA_CLAVICLE_YARKONI | unilateral | 125000\33088-6 |  | LN, 33088-6, Clavical length,Yarkoni 1985
OB-GYN Ultrasound Procedure Report | OB | GA(CRL)ASUM(2001) | none | calc | CALC_GA_CRL_ASUM2001 | unilateral | 125000\C12013-05 |  | 99PMSBLUS, C12013-05, CRL, ASUM 2001
OB-GYN Ultrasound Procedure Report | OB | GA(CRL)Hadlock | none | calc | CALC_GA_CRL_HADLOCK | unilateral | 125000\11910-7 |  | LN, 11910-7, CRL, Hadlock 1992
OB-GYN Ultrasound Procedure Report | OB | GA(CRL)Hansmann | none | calc | CALC_GA_CRL_HANSMANN | unilateral | 125000\33540-6 |  | LN, 33540-6, CRL, Hansmann 1986
OB-GYN Ultrasound Procedure Report | OB | GA(CRL)Jeanty | none | calc | CALC_GA_CRL_JEANTY | unilateral | 125000\11917-2 |  | LN, 11917-2, CRL, Jeanty 1984
OB-GYN Ultrasound Procedure Report | OB | GA(CRL)JSUM | none | calc | CALC_GA_CRL_JSUM | unilateral | 125000\C12013-22 |  | 99PMSBLUS, C12013-22, CRL, JSUM 2001
OB-GYN Ultrasound Procedure Report | OB | GA(CRL)Osaka | none | calc | CALC_GA_CRL_OSAKA | unilateral | 125000\C12013-32 |  | 99PMSBLUS, C12013-32, CRL, Osaka 1988
OB-GYN Ultrasound Procedure Report | OB | GA(CRL)Rempen | none | calc | CALC_GA_CRL_REMPEN | unilateral | 125000\33094-4 |  | LN, 33094-4, CRL, Rempen 1991
OB-GYN Ultrasound Procedure Report | OB | GA(CRL)Robinson | none | calc | CALC_GA_CRL_ROBINSON | unilateral | 125000\11914-9 |  | LN, 11914-9, CRL, Robinson 1975
OB-GYN Ultrasound Procedure Report | OB | GA(CRL)Robinson&Flem | none | calc | CALC_GA_CRL_ROBINSONFLEMING | unilateral | 125000\C12013-19 |  | 99PMSBLUS, C12013-19, CRL Robinson Fleming 1975
OB-GYN Ultrasound Procedure Report | OB | GA(CRL)Shinozuka | none | calc | CALC_GA_CRL_SHINOZUKA | unilateral | 125000\33095-1 |  | LN, 33095-1, CRL, Shinozuka 1996
OB-GYN Ultrasound Procedure Report | OB | GA(CRL)Tokyo | none | calc | CALC_GA_CRL_TOKYO | unilateral | 125000\C12013-27 |  | 99PMSBLUS, C12013-27, CRL, Tokyo 1989
OB-GYN Ultrasound Procedure Report | OB | GA(CRL)Verburg | none | calc | CALC_GA_CRL_VERBURG | unilateral | 125000\C12013-36 |  | 99PMSBLUS, C12013-36, CRL, Verburg 2008
OB-GYN Ultrasound Procedure Report | OB | GA(EFW)Shinozuka | none | calc | CALC_GA_EFW_SHINOZUKA | unilateral | 125000\C12013-20 |  | 99PMSBLUS, C12013-20, EFW, Shinozuka 2000
OB-GYN Ultrasound Procedure Report | OB | GA(FL)ASUM(2001) | none | calc | CALC_GA_FL_ASUM2001 | unilateral | 125000\C12013-06 |  | 99PMSBLUS, C12013-06, FL, ASUM 2001
OB-GYN Ultrasound Procedure Report | OB | GA(FL)Chitty | none | calc | CALC_GA_FL_CHITTY | unilateral | 125000\33098-5 |  | LN, 33098-5, FL, Chitty 1997
OB-GYN Ultrasound Procedure Report | OB | GA(FL)Hadlock | none | calc | CALC_GA_FL_HADLOCK | unilateral | 125000\11920-6 |  | LN, 11920-6, FL, Hadlock 1984
OB-GYN Ultrasound Procedure Report | OB | GA(FL)Hansmann | none | calc | CALC_GA_FL_HANSMANN | unilateral | 125000\33541-4 |  | LN, 33541-4, FL, Hansmann 1986
OB-GYN Ultrasound Procedure Report | OB | GA(FL)Jeanty | none | calc | CALC_GA_FL_JEANTY | unilateral | 125000\11923-0 |  | LN, 11923-0, FL, Jeanty 1984
OB-GYN Ultrasound Procedure Report | OB | GA(FL)Merz(1991) | none | calc | CALC_GA_FL_MERZ | unilateral | 125000\C12013-24 |  | 99PMSBLUS, C12013-24, FL, Merz 1991
OB-GYN Ultrasound Procedure Report | OB | GA(FL)Osaka | none | calc | CALC_GA_FL_OSAKA | unilateral | 125000\C12013-33 |  | 99PMSBLUS, C12013-33, FL, Osaka 1988
OB-GYN Ultrasound Procedure Report | OB | GA(FL)Shinozuka | none | calc | CALC_GA_FL_SHINOZUKA | unilateral | 125000\33102-5 |  | LN, 33102-5, FL, Shinozuka 1996
OB-GYN Ultrasound Procedure Report | OB | GA(FL)Tokyo | none | calc | CALC_GA_FL_TOKYO | unilateral | 125000\C12013-28 |  | 99PMSBLUS, C12013-28, FL, Tokyo 1989
OB-GYN Ultrasound Procedure Report | OB | GA(FL)Verburg | none | calc | CALC_GA_FL_VERBURG | unilateral | 125000\C12013-45 |  | 99PMSBLUS, C12013-45, FL, Verburg 2007
OB-GYN Ultrasound Procedure Report | OB | GA(FTA)Osaka | none | calc | CALC_GA_FTA_OSAKA | unilateral | 125000\C12013-35 |  | 99PMSBLUS, C12013-35, Fetal Trunk Cross-Sectional Area, Osaka 1988
OB-GYN Ultrasound Procedure Report | OB | GA(GSD)Rempen | none | calc | CALC_GA_GSD_REMPEN | unilateral | 125000\11929-7 |  | LN, 11929-7, GS, Rempen 1991
OB-GYN Ultrasound Procedure Report | OB | GA(GSD)Tokyo | none | calc | CALC_GA_GSD_TOKYO | unilateral | 125000\C12013-29 |  | 99PMSBLUS, C12013-29, GS, Tokyo 1989
OB-GYN Ultrasound Procedure Report | OB | GA(HC)ASUM(2001) | none | calc | CALC_GA_HC_ASUM2001 | unilateral | 125000\C12013-07 |  | 99PMSBLUS, C12013-07, HC, ASUM 2001
OB-GYN Ultrasound Procedure Report | OB | GA(HC)Chitty(Derived) | none | calc | CALC_GA_HC_CHITTY_DERIVED | unilateral | 125000\33111-6 |  | LN, 33111-6, HC derived, Chitty 1997
OB-GYN Ultrasound Procedure Report | OB | GA(HC)Chitty(Measured) | none | calc | CALC_GA_HC_CHITTY_MEASURED | unilateral | 125000\33110-8 |  | LN, 33110-8, HC measured, Chitty 1997
OB-GYN Ultrasound Procedure Report | OB | GA(HC)Hadlock | none | calc | CALC_GA_HC_HADLOCK | unilateral | 125000\11932-1 |  | LN, 11932-1, HC, Hadlock 1984
OB-GYN Ultrasound Procedure Report | OB | GA(HC)Hansmann | none | calc | CALC_GA_HC_HANSMANN | unilateral | 125000\33543-0 |  | LN, 33543-0, HC, Hansmann 1986
OB-GYN Ultrasound Procedure Report | OB | GA(HC)Merz(1991) | none | calc | CALC_GA_HC_MERZ | unilateral | 125000\C12013-25 |  | 99PMSBLUS, C12013-25, HC, Merz 1991
OB-GYN Ultrasound Procedure Report | OB | GA(HC)Verburg | none | calc | CALC_GA_HC_VERBURG | unilateral | 125000\C12013-44 |  | 99PMSBLUS, C12013-44, HC, Verburg 2007
OB-GYN Ultrasound Procedure Report | OB | GA(Humerus)ASUM(2001) | none | calc | CALC_GA_HUMERUS_ASUM2001 | unilateral | 125000\C12013-09 |  | 99PMSBLUS, C12013-09, Humerus, ASUM 2001
OB-GYN Ultrasound Procedure Report | OB | GA(Humerus)Jeanty | none | calc | CALC_GA_HL_JEANTY | unilateral | 125000\11936-2 |  | LN, 11936-2, Humerus, Jeanty 1984
OB-GYN Ultrasound Procedure Report | OB | GA(Humerus)Osaka | none | calc | CALC_GA_HL_OSAKA | unilateral | 125000\C12013-34 |  | 99PMSBLUS, C12013-34, Humerus Length, Osaka 1988
OB-GYN Ultrasound Procedure Report | OB | GA(LMP) | none | calc | CALC_GA_LMP | unilateral | 125000\11885-1 |  | LN, 11885-1, Gestational Age by LMP
OB-GYN Ultrasound Procedure Report | OB | GA(MSD)Hansmann | none | calc | CALC_GA_MSD_HANSMANN | unilateral | 125000\33106-6 |  | LN, 33106-6, GS, Hansmann 1982
OB-GYN Ultrasound Procedure Report | OB | GA(MSD)Hellman | none | calc | CALC_GA_MSD_HELLMAN | unilateral | 125000\11928-9 |  | LN, 11928-9, GS, Hellman 1969
OB-GYN Ultrasound Procedure Report | OB | GA(MSD)Nyberg | none | calc | CALC_GA_MSD_NYBERG | unilateral | 125000\33107-4 |  | LN, 33107-4, GS, Nyberg 1992
OB-GYN Ultrasound Procedure Report | OB | GA(OFD)ASUM(2001) | none | calc | CALC_GA_OFD_ASUM2001 | unilateral | 125000\C12013-11 |  | 99PMSBLUS, C12013-11, OFD, ASUM 2001
OB-GYN Ultrasound Procedure Report | OB | GA(OFD)Hansmann | none | calc | CALC_GA_OFD_HANSMANN | unilateral | 125000\33120-7 |  | LN, 33120-7, OFD, Hansmann 1986
OB-GYN Ultrasound Procedure Report | OB | GA(OOD)Jeanty | none | calc | CALC_GA_OOD_JEANTY | unilateral | 125000\C12013-12 |  | 99PMSBLUS, C12013-12, OOD, Jeanty 1984
OB-GYN Ultrasound Procedure Report | OB | GA(SL)Tokyo | none | calc | CALC_GA_SL_TOKYO | unilateral | 125000\33127-2 |  | LN, 33127-2, Spine Length, Tokyo, 1989
OB-GYN Ultrasound Procedure Report | OB | GA(TAD)Hansmann | none | calc | CALC_GA_TAD_HANSMANN | unilateral | 125000\C12013-41 |  | 99PMSBLUS, C12013-41, TAD, Hansmann 1986
OB-GYN Ultrasound Procedure Report | OB | GA(ThC)Nimrod | none | calc | CALC_GA_TC_NIMROD | unilateral | 125000\C12013-23 |  | 99PMSBLUS, C12013-23, TC, Nimrod 1986
OB-GYN Ultrasound Procedure Report | OB | GA(Tibia)Jeanty | none | calc | CALC_GA_TL_JEANTY | unilateral | 125000\11941-2 |  | LN, 11941-2, Tibia, Jeanty 1984
OB-GYN Ultrasound Procedure Report | OB | GA(TTD)Hansmann | none | calc | CALC_GA_TTD_HANSMANN | unilateral | 125000\C12013-18 |  | 99PMSBLUS, C12013-18, Transverse Trunk Diameter Hansmann 1986
OB-GYN Ultrasound Procedure Report | OB | GA(Ulna)Jeanty | none | calc | CALC_GA_UL_JEANTY | unilateral | 125000\11944-6 |  | LN, 11944-6, Ulna, Jeanty 1984
OB-GYN Ultrasound Procedure Report | OB | Gross Body Movement | none | meas | #fet_move | n/a | 125000\11631-9 |  | LN, 11631-9, Gross Body Movement
OB-GYN Ultrasound Procedure Report | OB | GSD1 | Length | meas | MEAS_GSD1 | unilateral | 125000\11850-5 |  | LN, 11850-5, Gestational Sac Diameter |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | GSD2 | Length | meas | MEAS_GSD2 | unilateral | 125000\11850-5 |  | LN, 11850-5, Gestational Sac Diameter |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | GSD3 | Length | meas | MEAS_GSD3 | unilateral | 125000\11850-5 |  | LN, 11850-5, Gestational Sac Diameter |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | HC | Circumference | meas | MEAS_HC_TRACED | unilateral | 125000\11984-2 |  | LN, 11984-2, Head Circumference |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | HC(c) | none | calc | CALC_HC | unilateral | 125000\11984-2 |  | LN, 11984-2, Head Circumference |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | HC/AC | none | calc | CALC_HC_AC | unilateral | 125000\11947-9 |  | LN, 11947-9, HC/AC |  | UCUM, 1, ratio
OB-GYN Ultrasound Procedure Report | OB | Head Angle | Alpha Angle | meas | m_head_angle | unilateral | 125000\C12019-06 |  | 99PMSBLUS, C12019-06, Head Angle |  | UCUM, deg, deg
OB-GYN Ultrasound Procedure Report | OB | Heart Axis Angle | Alpha Angle | meas | m_heart_axis_angle | unilateral | 125000\C12005-17 |  | 99PMSBLUS, C12005-17, Heart Axis Angle |  | UCUM, deg, deg
OB-GYN Ultrasound Procedure Report | OB | Hemisphere | Length | meas | m_Hem_width | unilateral | 125000\12170-7 |  | LN, 12170-7, Width of Hemisphere |  | UCUM, mm, Millimeter
OB-GYN Ultrasound Procedure Report | OB | HrtC | Circumference | meas | MEAS_HRT_CIRC | unilateral | 125000\C12005-03 |  | 99PMSBLUS, C12005-03, Heart Circumference |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | HrtC/ThC | none | calc | CALC_HRTC_TC | unilateral | 125000\C12004-01 |  | 99PMSBLUS, C12004-01, HrtC/TC |  | UCUM, 1, ratio
OB-GYN Ultrasound Procedure Report | OB | Humerus | Length | meas | MEAS_HL | unilateral | 125000\11966-9 |  | LN, 11966-9, Humerus length |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Iliac Angle | Alpha Angle | meas | m_iliac_angle | unilateral | 125000\C12005-15 |  | 99PMSBLUS, C12005-15, Iliac Angle |  | UCUM, deg, deg
OB-GYN Ultrasound Procedure Report | OB | Inferior Facial Angle | Alpha Angle | meas | m_inferior_facial_angle | unilateral | 125000\C12005-13 |  | 99PMSBLUS, C12005-13, Inferior Facial Angle |  | UCUM, deg, deg
OB-GYN Ultrasound Procedure Report | OB | IOD | Length | meas | MEAS_IOD | unilateral | 125000\33070-4 |  | LN, 33070-4, Inner Orbital Diameter |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | L Ovary AI | Doppler Slope | meas | m_LOv_AS | unilateral | 125000\20167-3\G-A101\T-46980 | SRT, T-46980, Ovarian Artery | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [8 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Ovary AT | Time | meas | m_LOv_AT | unilateral | 125000\20168-1\G-A101\T-46980 | SRT, T-46980, Ovarian Artery | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [8 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Ovary D/S | none | calc | c_LOv_DS | unilateral | 125000\C12121-01\G-A101\T-46980 | SRT, T-46980, Ovarian Artery | 99PMSBLUS, C12121-01, Diastolic to Systolic Velocity Ratio |  | UCUM, 1, ratio | [8 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Ovary EDV | Signed End Diastolic Vel | meas | m_LOv_EDV | unilateral | 125000\11653-3\G-A101\T-46980 | SRT, T-46980, Ovarian Artery | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [8 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Ovary MDV | Signed Minimum Velocity | meas | m_LOv_MDV | unilateral | 125000\11665-7\G-A101\T-46980 | SRT, T-46980, Ovarian Artery | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [8 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Ovary PI | none | calc | c_LOv_PI | unilateral | 125000\12008-9\G-A101\T-46980 | SRT, T-46980, Ovarian Artery | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [8 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Ovary PSV | Signed Maximum Velocity | meas | m_LOv_PSV | unilateral | 125000\11726-7\G-A101\T-46980 | SRT, T-46980, Ovarian Artery | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [8 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Ovary RI | none | calc | c_LOv_RI | unilateral | 125000\12023-8\G-A101\T-46980 | SRT, T-46980, Ovarian Artery | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [8 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Ovary S/D | none | calc | c_LOv_SD | unilateral | 125000\12144-2\G-A101\T-46980 | SRT, T-46980, Ovarian Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [8 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Ovary TAMV | Time Average Mean Velocity | meas | m_LOv_TAMV | unilateral | 125000\20352-1\G-A101\T-46980 | SRT, T-46980, Ovarian Artery | LN, 20352-1, Time Averaged Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [8 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Ovary TAPV | Signed Time Average Peak Vel | meas | m_LOv_TAPV | unilateral | 125000\11692-1\G-A101\T-46980 | SRT, T-46980, Ovarian Artery | LN, 11692-1, Time Averaged Peak Velocity |  | UCUM, cm/s, Centimeter Per Second | [8 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Ovary VTI | Velocity Time Integral | meas | m_LOv_VTI | unilateral | 125000\20354-7\G-A101\T-46980 | SRT, T-46980, Ovarian Artery | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [8 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Uterine Art PI | none | calc | c_Lt_UterArt_PI | unilateral | 125000\12008-9\G-A101\T-46820 | SRT, T-46820, Uterine Artery | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [8 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Uterine Art RI | none | calc | c_Lt_UterArt_RI | unilateral | 125000\12023-8\G-A101\T-46820 | SRT, T-46820, Uterine Artery | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [8 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | L Uterine Art S/D | none | calc | c_Lt_UterArt_SD | unilateral | 125000\12144-2\G-A101\T-46820 | SRT, T-46820, Uterine Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [8 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | Lat Vent Anterior | Length | meas | m_LAT_Va_width | unilateral | 125000\33197-5 |  | LN, 33197-5, Anterior Horn Lateral ventricular width |  | UCUM, mm, Millimeter
OB-GYN Ultrasound Procedure Report | OB | Lat Vent Post/Hem | none | calc | c_OB_LatVp_Hem_RATIO | unilateral | 125000\C12007-03 |  | 99PMSBLUS, C12007-03, Ratio of Posterior Horn Lateral ventricular width/Hemisphere |  | UCUM, 1, ratio
OB-GYN Ultrasound Procedure Report | OB | Lat Vent Posterior | Length | meas | m_LAT_Vp_width | unilateral | 125000\33196-7 |  | LN, 33196-7, Posterior Horn Lateral ventricular width |  | UCUM, mm, Millimeter
OB-GYN Ultrasound Procedure Report | OB | Lat Ventricle | Length | meas | MEAS_LAT_V | unilateral | 125000\12171-5 |  | LN, 12171-5, Lateral Ventrical width |  | UCUM, mm, Millimeter
OB-GYN Ultrasound Procedure Report | OB | LOH | Length | meas | MEAS_LEFT_OVARY_HEIGHT | unilateral | 125000\11857-0 |  | LN, 11857-0, Left Ovary Height |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | LOL | Length | meas | MEAS_LEFT_OVARY_LENGTH | unilateral | 125000\11840-6 |  | LN, 11840-6, Left Ovary Length |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | LOV | none | calc | CALC_LEFT_OVARY_VOLUME | unilateral | 125000\12164-0 |  | LN, 12164-0, Left Ovary Volume |  | UCUM, cm3, Cubic Centimeter
OB-GYN Ultrasound Procedure Report | OB | LOW | Length | meas | MEAS_LEFT_OVARY_WIDTH | unilateral | 125000\11829-9 |  | LN, 11829-9, Left Ovary Width |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | M Phalanx 5 | Length | meas | MEAS_M_PHALANX_5 | unilateral | 125000\C12005-04 |  | 99PMSBLUS, C12005-04, Length of Middle Phalanx of the Fifth Digit |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | MCA AI | Doppler Slope | meas | m_OB_MCA_AS | unilateral | 125000\ T-45600\20167-3 | SRT, T-45600, Middle Cerebral Artery | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square
OB-GYN Ultrasound Procedure Report | OB | MCA AT | Time | meas | m_OB_MCA_AT | unilateral | 125000\ T-45600\20168-1 | SRT, T-45600, Middle Cerebral Artery | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds
OB-GYN Ultrasound Procedure Report | OB | MCA D/S | none | calc | c_OB_MCA_DS | unilateral | 125000\ T-45600\C12121-01 | SRT, T-45600, Middle Cerebral Artery | 99PMSBLUS, C12121-01, Diastolic to Systolic Velocity Ratio |  | UCUM, 1, ratio
OB-GYN Ultrasound Procedure Report | OB | MCA EDV | Signed End Diastolic Vel | meas | m_OB_MCA_EDV | unilateral | 125000\ T-45600\11653-3 | SRT, T-45600, Middle Cerebral Artery | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second
OB-GYN Ultrasound Procedure Report | OB | MCA MDV | Signed Minimum Velocity | meas | m_OB_MCA_MDV | unilateral | 125000\ T-45600\11665-7 | SRT, T-45600, Middle Cerebral Artery | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second
OB-GYN Ultrasound Procedure Report | OB | MCA PI | none | calc | c_OB_MCA_PI | unilateral | 125000\ T-45600\12008-9 | SRT, T-45600, Middle Cerebral Artery | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio
OB-GYN Ultrasound Procedure Report | OB | MCA PSV | Signed Maximum Velocity | meas | m_OB_MCA_PSV | unilateral | 125000\ T-45600\11726-7 | SRT, T-45600, Middle Cerebral Artery | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second
OB-GYN Ultrasound Procedure Report | OB | MCA RI | none | calc | c_OB_MCA_RI | unilateral | 125000\ T-45600\12023-8 | SRT, T-45600, Middle Cerebral Artery | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio
OB-GYN Ultrasound Procedure Report | OB | MCA S/D | none | calc | c_OB_MCA_SD | unilateral | 125000\ T-45600\12144-2 | SRT, T-45600, Middle Cerebral Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio
OB-GYN Ultrasound Procedure Report | OB | MCA TAMV | Time Average Mean Velocity | meas | m_OB_MCA_TAMV | unilateral | 125000\ T-45600\20352-1 | SRT, T-45600, Middle Cerebral Artery | LN, 20352-1, Time Averaged Mean Velocity |  | UCUM, cm/s, Centimeter Per Second
OB-GYN Ultrasound Procedure Report | OB | MCA TAPV | Signed Time Average Peak Vel | meas | m_OB_MCA_TAPV | unilateral | 125000\ T-45600\11692-1 | SRT, T-45600, Middle Cerebral Artery | LN, 11692-1, Time Averaged Peak Velocity |  | UCUM, cm/s, Centimeter Per Second
OB-GYN Ultrasound Procedure Report | OB | MCA VTI | Velocity Time Integral | meas | m_OB_MCA_VTI | unilateral | 125000\ T-45600\20354-7 | SRT, T-45600, Middle Cerebral Artery | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | MDV | Signed Minimum Velocity | meas | m_Lt_UterArt_MDV | unilateral | 125000\11665-7\G-A101\T-46820 | SRT, T-46820, Uterine Artery | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [8 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | MDV | Signed Minimum Velocity | meas | m_Rt_UterArt_MDV | unilateral | 125000\11665-7\G-A100\T-46820 | SRT, T-46820, Uterine Artery | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [8 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | MSD | none | calc | CALC_MSD | unilateral | 125000\C12009-04 |  | 99PMSBLUS, C12009-04, Mean Sac Diameter |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Nuchal | Length | meas | MEAS_NUCHAL_THICKNESS | unilateral | 125000\12146-7 |  | LN, 12146-7, Nuchal Fold thickness |  | UCUM, mm, Millimeter
OB-GYN Ultrasound Procedure Report | OB | Number of follicles in left ovary | none | meas | #leftOvFlcNumber | n/a | 125000\11879-4 |  | LN, 11879-4, Number of follicles in left ovary
OB-GYN Ultrasound Procedure Report | OB | Number of follicles in right ovary | none | meas | #rightOvFlcNumber | n/a | 125000\11880-2 |  | LN, 11880-2, Number of follicles in right ovary
OB-GYN Ultrasound Procedure Report | OB | OFD | Length | meas | MEAS_OFD | unilateral | 125000\11851-3 |  | LN, 11851-3, Occipital-Frontal Diameter |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | OOD | Length | meas | MEAS_OOD | unilateral | 125000\11629-3 |  | LN, 11629-3, Outer Orbital Diameter |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Orbit 1 | Length | meas | MEAS_ORBIT1 | unilateral | 125000\C12007-01 |  | 99PMSBLUS, C12007-01, Diameter of First Orbit |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Orbit 2 | Length | meas | MEAS_ORBIT2 | unilateral | 125000\C12007-02 |  | 99PMSBLUS, C12007-02, Diameter of Second Orbit |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | PSV | Signed Maximum Velocity | meas | m_Lt_UterArt_PSV | unilateral | 125000\11726-7\G-A101\T-46820 | SRT, T-46820, Uterine Artery | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [8 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | PSV | Signed Maximum Velocity | meas | m_Rt_UterArt_PSV | unilateral | 125000\11726-7\G-A100\T-46820 | SRT, T-46820, Uterine Artery | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [8 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | PV Bladder Height | Length | meas | MEAS_GYN_PVBLADDER_HEIGHT | unilateral | 125000\C12011-07 |  | 99PMSBLUS, C12011-07, Post Void Bladder Height |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | PV Bladder Length | Length | meas | MEAS_GYN_PVBLADDER_LENGTH | unilateral | 125000\C12011-05 |  | 99PMSBLUS, C12011-05, Post Void Bladder Length |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | PV Bladder Vol | none | calc | CALC_GYN_PVBLADDER_VOL | unilateral | 125000\C12011-08 |  | 99PMSBLUS, C12011-08, Post Void Bladder Volume |  | UCUM, cm3, Cubic Centimeter
OB-GYN Ultrasound Procedure Report | OB | PV Bladder Width | Length | meas | MEAS_GYN_PVBLADDER_WIDTH | unilateral | 125000\C12011-06 |  | 99PMSBLUS, C12011-06, Post Void Bladder Width |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Q1 | Length | meas | MEAS_QUAD1 | unilateral | 125000\11624-4 |  | LN, 11624-4, First Quadrant Diameter |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Q2 | Length | meas | MEAS_QUAD2 | unilateral | 125000\11626-9 |  | LN, 11626-9, Second Quadrant Diameter |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Q3 | Length | meas | MEAS_QUAD3 | unilateral | 125000\11625-1 |  | LN, 11625-1, Third Quadrant Diameter |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Q4 | Length | meas | MEAS_QUAD4 | unilateral | 125000\11623-6 |  | LN, 11623-6, Fourth Quadrant Diameter |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | R Ovary AI | Doppler Slope | meas | m_ROv_AS | unilateral | 125000\20167-3\G-A100\T-46980 | SRT, T-46980, Ovarian Artery | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [8 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Ovary AT | Time | meas | m_ROv_AT | unilateral | 125000\20168-1\G-A100\T-46980 | SRT, T-46980, Ovarian Artery | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [8 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Ovary D/S | none | calc | c_ROv_DS | unilateral | 125000\C12121-01\G-A100\T-46980 | SRT, T-46980, Ovarian Artery | 99PMSBLUS, C12121-01, Diastolic to Systolic Velocity Ratio |  | UCUM, 1, ratio | [8 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Ovary EDV | Signed End Diastolic Vel | meas | m_ROv_EDV | unilateral | 125000\11653-3\G-A100\T-46980 | SRT, T-46980, Ovarian Artery | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [8 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Ovary MDV | Signed Minimum Velocity | meas | m_ROv_MDV | unilateral | 125000\11665-7\G-A100\T-46980 | SRT, T-46980, Ovarian Artery | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [8 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Ovary PI | none | calc | c_ROv_PI | unilateral | 125000\12008-9\G-A100\T-46980 | SRT, T-46980, Ovarian Artery | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [8 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Ovary PSV | Signed Maximum Velocity | meas | m_ROv_PSV | unilateral | 125000\11726-7\G-A100\T-46980 | SRT, T-46980, Ovarian Artery | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [8 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Ovary RI | none | calc | c_ROv_RI | unilateral | 125000\12023-8\G-A100\T-46980 | SRT, T-46980, Ovarian Artery | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [8 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Ovary S/D | none | calc | c_ROv_SD | unilateral | 125000\12144-2\G-A100\T-46980 | SRT, T-46980, Ovarian Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [8 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Ovary TAMV | Time Average Mean Velocity | meas | m_ROv_TAMV | unilateral | 125000\20352-1\G-A100\T-46980 | SRT, T-46980, Ovarian Artery | LN, 20352-1, Time Averaged Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [8 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Ovary TAPV | Signed Time Average Peak Vel | meas | m_ROv_TAPV | unilateral | 125000\11692-1\G-A100\T-46980 | SRT, T-46980, Ovarian Artery | LN, 11692-1, Time Averaged Peak Velocity |  | UCUM, cm/s, Centimeter Per Second | [8 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Ovary VTI | Velocity Time Integral | meas | m_ROv_VTI | unilateral | 125000\20354-7\G-A100\T-46980 | SRT, T-46980, Ovarian Artery | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [8 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Uterine Art PI | none | calc | c_Rt_UterArt_PI | unilateral | 125000\12008-9\G-A100\T-46820 | SRT, T-46820, Uterine Artery | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [8 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Uterine Art RI | none | calc | c_Rt_UterArt_RI | unilateral | 125000\12023-8\G-A100\T-46820 | SRT, T-46820, Uterine Artery | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [8 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | R Uterine Art S/D | none | calc | c_Rt_UterArt_SD | unilateral | 125000\12144-2\G-A100\T-46820 | SRT, T-46820, Uterine Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [8 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | Renal AP | Length | meas | MEAS_RENAL_AP | unilateral | 125000\C12005-05 |  | 99PMSBLUS, C12005-05, Renal Width |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Renal L | Length | meas | MEAS_RENAL_LENGTH | unilateral | 125000\C12005-06 |  | 99PMSBLUS, C12005-06, Renal Length |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | RL | Length | meas | MEAS_RL | unilateral | 125000\11967-7 |  | LN, 11967-7, Radius length |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | ROH | Length | meas | MEAS_RIGHT_OVARY_HEIGHT | unilateral | 125000\11858-8 |  | LN, 11858-8, Right Ovary Height |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | ROL | Length | meas | MEAS_RIGHT_OVARY_LENGTH | unilateral | 125000\11841-4 |  | LN, 11841-4, Right Ovary Length |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | ROV | none | calc | CALC_RIGHT_OVARY_VOLUME | unilateral | 125000\12165-7 |  | LN, 12165-7, Right Ovary Volume |  | UCUM, cm3, Cubic Centimeter
OB-GYN Ultrasound Procedure Report | OB | ROW | Length | meas | MEAS_RIGHT_OVARY_WIDTH | unilateral | 125000\11830-7 |  | LN, 11830-7, Right Ovary Width |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | SL | Length | meas | MEAS_SL | unilateral | 125000\33071-2 |  | LN, 33071-2, Spine Length |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | TAD | Length | meas | MEAS_AD_TRV | unilateral | 125000\11862-0 |  | LN, 11862-0, Tranverse Abdominal Diameter |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | TAPV | Signed Time Average Peak Vel | meas | m_Lt_UterArt_TAPV | unilateral | 125000\11692-1\G-A101\T-46820 | SRT, T-46820, Uterine Artery | LN, 11692-1, Time Averaged Peak Velocity |  | UCUM, cm/s, Centimeter Per Second | [8 empty] | SRT, G-A101, Left
OB-GYN Ultrasound Procedure Report | OB | TAPV | Signed Time Average Peak Vel | meas | m_Rt_UterArt_TAPV | unilateral | 125000\11692-1\G-A100\T-46820 | SRT, T-46820, Uterine Artery | LN, 11692-1, Time Averaged Peak Velocity |  | UCUM, cm/s, Centimeter Per Second | [8 empty] | SRT, G-A100, Right
OB-GYN Ultrasound Procedure Report | OB | ThC | Circumference | meas | MEAS_TC_TRACED | unilateral | 125000\11988-3 |  | LN, 11988-3, Thoracic Circumference |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | ThC(c) | none | calc | CALC_TC | unilateral | 125000\11988-3 |  | LN, 11988-3, Thoracic Circumference |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Tibia | Length | meas | MEAS_TL | unilateral | 125000\11968-5 |  | LN, 11968-5, Tibia length |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | TTD | Length | meas | MEAS_TTD | unilateral | 125000\C12005-08 |  | 99PMSBLUS, C12005-08, Transverse Trunk Diameter |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | TThD | Length | meas | MEAS_TD_TRV | unilateral | 125000\11864-6 |  | LN, 11864-6, Transverse Thoracic Diameter |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Ulna | Length | meas | MEAS_UL | unilateral | 125000\11969-3 |  | LN, 11969-3, Ulna length |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Umbilical AI | Doppler Slope | meas | m_Umb_AS | unilateral | 125000\20167-3\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square
OB-GYN Ultrasound Procedure Report | OB | Umbilical AT | Time | meas | m_Umb_AT | unilateral | 125000\20168-1\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds
OB-GYN Ultrasound Procedure Report | OB | Umbilical D/S | none | calc | c_Umb_DS | unilateral | 125000\C12121-01\T-F1810 | SRT, T-F1810, Umbilical Artery | 99PMSBLUS, C12121-01, Diastolic to Systolic Velocity Ratio |  | UCUM, 1, ratio
OB-GYN Ultrasound Procedure Report | OB | Umbilical EDV | Signed End Diastolic Vel | meas | m_Umb_EDV | unilateral | 125000\11653-3\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second
OB-GYN Ultrasound Procedure Report | OB | Umbilical MDV | Signed Minimum Velocity | meas | m_Umb_MDV | unilateral | 125000\11665-7\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second
OB-GYN Ultrasound Procedure Report | OB | Umbilical PI | none | calc | c_Umb_PI | unilateral | 125000\12008-9\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio
OB-GYN Ultrasound Procedure Report | OB | Umbilical PSV | Signed Maximum Velocity | meas | m_Umb_PSV | unilateral | 125000\11726-7\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second
OB-GYN Ultrasound Procedure Report | OB | Umbilical RI | none | calc | c_Umb_RI | unilateral | 125000\12023-8\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio
OB-GYN Ultrasound Procedure Report | OB | Umbilical S/D | none | calc | c_Umb_SD | unilateral | 125000\12144-2\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio
OB-GYN Ultrasound Procedure Report | OB | Umbilical TAMV | Time Average Mean Velocity | meas | m_Umb_TAMV | unilateral | 125000\20352-1\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 20352-1, Time Averaged Mean Velocity |  | UCUM, cm/s, Centimeter Per Second
OB-GYN Ultrasound Procedure Report | OB | Umbilical TAPV | Signed Time Average Peak Vel | meas | m_Umb_TAPV | unilateral | 125000\11692-1\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 11692-1, Time Averaged Peak Velocity |  | UCUM, cm/s, Centimeter Per Second
OB-GYN Ultrasound Procedure Report | OB | Umbilical VTI | Velocity Time Integral | meas | m_Umb_VTI | unilateral | 125000\20354-7\T-F1810 | SRT, T-F1810, Umbilical Artery | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Uterine AI | Doppler Slope | meas | m_Uterine_AS | unilateral | 125000\20167-3\T-46820 | SRT, T-46820, Uterine Artery | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square
OB-GYN Ultrasound Procedure Report | OB | Uterine AT | Time | meas | m_Uterine_AT | unilateral | 125000\20168-1\T-46820 | SRT, T-46820, Uterine Artery | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds
OB-GYN Ultrasound Procedure Report | OB | Uterine D/S | none | calc | c_Uterine_DS | unilateral | 125000\C12121-01\T-46820 | SRT, T-46820, Uterine Artery | 99PMSBLUS, C12121-01, Diastolic to Systolic Velocity Ratio |  | UCUM, 1, ratio
OB-GYN Ultrasound Procedure Report | OB | Uterine EDV | Signed End Diastolic Vel | meas | m_Uterine_EDV | unilateral | 125000\11653-3\T-46820 | SRT, T-46820, Uterine Artery | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second
OB-GYN Ultrasound Procedure Report | OB | Uterine MDV | Signed Minimum Velocity | meas | m_Uterine_MDV | unilateral | 125000\11665-7\T-46820 | SRT, T-46820, Uterine Artery | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second
OB-GYN Ultrasound Procedure Report | OB | Uterine PI | none | calc | c_Uterine_PI | unilateral | 125000\12008-9\T-46820 | SRT, T-46820, Uterine Artery | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio
OB-GYN Ultrasound Procedure Report | OB | Uterine PSV | Signed Maximum Velocity | meas | m_Uterine_PSV | unilateral | 125000\11726-7\T-46820 | SRT, T-46820, Uterine Artery | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second
OB-GYN Ultrasound Procedure Report | OB | Uterine RI | none | calc | c_Uterine_RI | unilateral | 125000\12023-8\T-46820 | SRT, T-46820, Uterine Artery | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio
OB-GYN Ultrasound Procedure Report | OB | Uterine S/D | none | calc | c_Uterine_SD | unilateral | 125000\12144-2\T-46820 | SRT, T-46820, Uterine Artery | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio
OB-GYN Ultrasound Procedure Report | OB | Uterine TAMV | Time Average Mean Velocity | meas | m_Uterine_TAMV | unilateral | 125000\20352-1\T-46820 | SRT, T-46820, Uterine Artery | LN, 20352-1, Time Averaged Mean Velocity |  | UCUM, cm/s, Centimeter Per Second
OB-GYN Ultrasound Procedure Report | OB | Uterine TAPV | Signed Time Average Peak Vel | meas | m_Uterine_TAPV | unilateral | 125000\11692-1\T-46820 | SRT, T-46820, Uterine Artery | LN, 11692-1, Time Averaged Peak Velocity |  | UCUM, cm/s, Centimeter Per Second
OB-GYN Ultrasound Procedure Report | OB | Uterine VTI | Velocity Time Integral | meas | m_Uterine_VTI | unilateral | 125000\20354-7\T-46820 | SRT, T-46820, Uterine Artery | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | UTH | Length | meas | MEAS_UTERINE_HEIGHT | unilateral | 125000\11859-6 |  | LN, 11859-6, Uterus Height |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | UTL | Length | meas | MEAS_UTERINE_LENGTH | unilateral | 125000\11842-2 |  | LN, 11842-2, Uterus Length |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | UTV | none | calc | CALC_UTERINE_VOLUME | unilateral | 125000\33192-6 |  | LN, 33192-6, Uterus Volume |  | UCUM, cm3, Cubic Centimeter
OB-GYN Ultrasound Procedure Report | OB | UTW | Length | meas | MEAS_UTERINE_WIDTH | unilateral | 125000\11865-3 |  | LN, 11865-3, Uterus Width |  | UCUM, cm, Centimeter
OB-GYN Ultrasound Procedure Report | OB | Yolk Sac | Length | meas | MEAS_YOLKSAC | unilateral | 125000\11816-6 |  | LN, 11816-6, Yolk Sac length |  | UCUM, cm, Centimeter
Patient Characteristics | Patient Characteristics | Aborta | none |  | #aborta | n/a | 11612-9 |  |  | LN, 11612-9, Aborta
Patient Characteristics | Patient Characteristics | Body Surface Area | none |  | #BodySurArea | n/a | 2329305 |  |  | LN, 8277-6, Body Surface Area
Patient Characteristics | Patient Characteristics | Body Surface Area Formula | none |  | #BodySurAreaFormula | n/a | 2318652 |  |  | LN, 8248-4, Body Surface Area Formula
Patient Characteristics | Patient Characteristics | Diastolic Blood Pressure | none |  | #DiastBloodPres | n/a | F-008ED |  |  | SRT, F-008ED, Diastolic Blood Pressure
Patient Characteristics | Patient Characteristics | Ectopic Pregnancies | none |  | #ectopic | n/a | 33065-4 |  |  | LN, 33065-4, Ectopic Pregnancies
Patient Characteristics | Patient Characteristics | Gravida | none |  | #gravida | n/a | 11996-6 |  |  | LN, 11996-6, Gravida
Patient Characteristics | Patient Characteristics | Para | none |  | #para | n/a | 11977-6 |  |  | LN, 11977-6, Para
Patient Characteristics | Patient Characteristics | Patient Height | none |  | #height | n/a | 2338315 |  |  | LN, 8302-2, Patient Height
Patient Characteristics | Patient Characteristics | Patient ID | none |  | #patient_id | n/a | patient_id |  |  | 99PMSBLU, patient_id, Patient ID
Patient Characteristics | Patient Characteristics | Patient Information | none |  | #patient_info | n/a | patient_info |  |  | 99PMSBLU, patient_info, Patient Information
Patient Characteristics | Patient Characteristics | Patient Name | none |  | #patient_name | n/a | patient_name |  |  | 99PMSBLU, patient_name, Patient Name
Patient Characteristics | Patient Characteristics | Patient Weight | none |  | #weight | n/a | 29463-7 |  |  | LN, 29463-7, Patient Weight
Patient Characteristics | Patient Characteristics | Subject Age | none |  | #SubjectAge | n/a | 121033 |  |  | DCM, 121033, Subject Age
Patient Characteristics | Patient Characteristics | Subject Sex | none |  | #SubjectSex | n/a | 121032 |  |  | DCM, 121032, Subject Sex
Patient Characteristics | Patient Characteristics | Systolic Blood Pressure | none |  | #SystBloodPres | n/a | F-008EC |  |  | SRT, F-008EC, Systolic Blood Pressure
Pediatric Cardiac Ultrasound Report | Pediatric Echo | A Wave Amp | Length | meas | m_p_AWaveAmp | unilateral | 125195\59100-8\G-0394\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 59100-8, A-Wave Amplitude |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Accel Slope | Doppler Slope | meas | m_p_AI_AccelSlope_DopSlope | unilateral | 125195\20167-3\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20167-3, Acceleration Slope |  | UCUM, cm/s2, Centimeter Per Second Square |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Accel Time | Time | meas | m_p_AI_AccelSlope_DopTime | unilateral | 125195\20168-1\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Alias Vel | Unsigned Maximum Velocity | meas | m_p_AI_AliasVel | unilateral | 125195\59130-5\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 59130-5, Alias velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Decel Slope | Doppler Slope | meas | m_p_AI_DS_DopSlope | unilateral | 125195\20216-8\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20216-8, Deceleration Slope |  | UCUM, cm/s2, Centimeter Per Second Square |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Decel Time | Time | meas | m_p_AI_DS_DopTime | unilateral | 125195\20217-6\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20217-6, Deceleration Time |  | UCUM, sec, Seconds |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI End Dias Vel | Unsigned Maximum Velocity | meas | m_p_AI_EndDiVel_Vel | unilateral | 125195\11653-3\F-32011\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  | SRT, F-32011, End Diastole | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI ERO | none | calc | c_p_AI_ERO | unilateral | 125195\125216\G-0367\G-038E\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Flow Rate | none | calc | c_p_AI_FlowRate | unilateral | 125195\34141-2\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 34141-2, Peak Instantaneous Flow Rate |  | UCUM, ml/sec, ml/sec |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Fraction | none | calc | c_p_AI_Frac | unilateral | 125195\G-0367\G-0390\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-0390, Regurgitant Fraction |  | UCUM, %, Percent |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Max PG | Maximum PG | meas | m_p_AI_Vmax_MaxPG | unilateral | 125195\20247-3\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Mean PG | Mean PG | meas | m_p_AI_VTI_CardPGmean | unilateral | 125195\20256-4\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI P1/2t | Pressure Half Time | meas | m_p_AI_PHalfT_PHalftime | unilateral | 125195\20280-4\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20280-4, Pressure Half-Time |  | UCUM, ms, Millisecond |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Radius | Length | meas | m_p_AI_Radius | unilateral | 125195\59102-4\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 59102-4, Flow Radius |  | UCUM, cm, Centimeter |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Vmax | Unsigned Maximum Velocity | meas | m_p_AI_Vmax_Vmax | unilateral | 125195\20351-3\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Vmean | Unsigned Time Average Peak Vel | meas | m_p_AI_VTI_CardTAPV | unilateral | 125195\11692-1\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 11692-1, Time Averaged Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI Volume | none | calc | c_p_AI_Vol | unilateral | 125195\125216\33878-0\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 33878-0, Volume Flow |  | UCUM, ml, Milliliter |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AI VTI | Velocity Time Integral | meas | m_p_AI_VTI_VTI | unilateral | 125195\20354-7\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Ao Arch Diam | Length | meas | m_p_AoArch_Diam | unilateral | 125195\G-03A2\M-02550\T-42300 | SRT, T-42300, Aortic arch | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Ao Arch Dist Diam | Length | meas | m_p_AoArch_Dist_Diam | unilateral | 125195\G-03A2\G-A119\M-02550\T-42300 | SRT, T-42300, Aortic arch | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Ao Isthmus Diam | Length | meas | m_p_AoIsth_Diam | unilateral | 125195\G-03A2\M-02550\T-42310 | SRT, T-42310, Aortic isthmus | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Ao Sinus Diam | Length | meas | m_p_Ao_Sinus_Diam | unilateral | 125195\G-03A2\M-02550\T-42200 | SRT, T-42200, Structure Sinus of Valsalva | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Ao ST Jx Diam | Length | meas | m_p_ST_J_Diam | unilateral | 125195\G-03A2\M-02550\T-42102 | SRT, T-42102, Aortic sinotubular junction | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AoR Area | none | calc | c_p_AoR_Area | unilateral | 125195\G-038E\G-03A2\T-42110 | SRT, T-42110, Root of Aorta | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AoR Diam (2D) | Length | meas | m_p_AoR_Diam_2D | unilateral | 125195\G-03A2\M-02550\T-42110 | SRT, T-42110, Root of Aorta | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AoR Diam (MM) | Length | meas | m_p_AoR_Diam_MM | unilateral | 125195\G-0394\M-02550\T-42110 | SRT, T-42110, Root of Aorta | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AS Max PG | Maximum PG | meas | m_p_AS_Vmax_MaxPG | unilateral | 125195\59106-5\T-35400 | SRT, T-35400, Aortic Valve | LN, 59106-5, Stenosis Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AS Vmax | Unsigned Maximum Velocity | meas | m_p_AS_Vmax_Vmax | unilateral | 125195\59107-3\T-35400 | SRT, T-35400, Aortic Valve | LN, 59107-3, Stenosis Peak Velocity |  | UCUM, cm/s, Centimeter Per Second
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Asc Ao Diam | Length | meas | m_p_AscAorta_Diam | unilateral | 125195\G-03A2\M-02550\T-42100 | SRT, T-42100, Ascending aorta | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Asc Ao Max PG | Maximum PG | meas | m_p_ASC_Ao_MaxPG_MaxPG | unilateral | 125195\20247-3\T-42100 | SRT, T-42100, Ascending aorta | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Asc Ao Max PG (full) | none | calc | c_p_AscAo_MaxPG_FB | unilateral | 125195\125217\20247-3\T-42100 | SRT, T-42100, Ascending aorta | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury | [5 empty] | DCM, 125217, Full Bernoulli
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Asc Ao Mean PG | Mean PG | meas | m_p_ASC_Ao_PG | unilateral | 125195\20256-4\T-42100 | SRT, T-42100, Ascending aorta | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Asc Ao Mean PG (full) | none | calc | c_p_AscAo_MeanPG_FB | unilateral | 125195\125217\20256-4\T-42100 | SRT, T-42100, Ascending aorta | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury | [5 empty] | DCM, 125217, Full Bernoulli
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Asc Ao Vmax | Unsigned Maximum Velocity | meas | m_p_ASC_Ao_MaxPG_Vmax | unilateral | 125195\20351-3\T-42100 | SRT, T-42100, Ascending aorta | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ASD Diam | none | calc | c_p_ASD_Diam | unilateral | 125195\D4-31220\G-03A2\M-02550 | SRT, D4-31220, Atrial Septal Defect | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ASD Major | Length | meas | m_p_ASD_Major | unilateral | 125195\D4-31220\G-03A2\G-A193 | SRT, D4-31220, Atrial Septal Defect | SRT, G-A193, Major Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ASD Max PG | Maximum PG | meas | m_p_ASD_Vmax_MaxPG | unilateral | 125195\20247-3\D4-31220 | SRT, D4-31220, Atrial Septal Defect | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ASD Mean PG | Mean PG | meas | m_p_ASD_VTI_CardPGmean | unilateral | 125195\20256-4\D4-31220 | SRT, D4-31220, Atrial Septal Defect | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ASD Minor | Length | meas | m_p_ASD_Minor | unilateral | 125195\D4-31220\G-03A2\G-A194 | SRT, D4-31220, Atrial Septal Defect | SRT, G-A194, Minor Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ASD Vmax | Unsigned Maximum Velocity | meas | m_p_ASD_Vmax_Vmax | unilateral | 125195\20351-3\D4-31220 | SRT, D4-31220, Atrial Septal Defect | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ASD Vmean | Unsigned Time Average Peak Vel | meas | m_p_ASD_VTI_CardTAPV | unilateral | 125195\11692-1\D4-31220 | SRT, D4-31220, Atrial Septal Defect | LN, 11692-1, Time Averaged Peak Velocity |  | UCUM, cm/s, Centimeter Per Second
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ASD VTI | Velocity Time Integral | meas | m_p_ASD_VTI_VTI | unilateral | 125195\20354-7\D4-31220 | SRT, D4-31220, Atrial Septal Defect | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Accel Slope | Doppler Slope | meas | m_p_AV_AccelSlope_DopSlope | unilateral | 125195\20167-3\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20167-3, Acceleration Slope |  | UCUM, cm/s2, Centimeter Per Second Square |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Accel Time | Time | meas | m_p_AV_AccelSlope_DopTime | unilateral | 125195\20168-1\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Annul Diam | Length | meas | m_p_AV_Annul_Diam | unilateral | 125195\G-03A2\M-02550\T-3500E\T-35400 | SRT, T-35400, Aortic Valve | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter | SRT, T-3500E, Cardiac valve annulus | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Area | Area | meas | m_p_AV_Area | unilateral | 125195\G-03A2\G-A166\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-A166, Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Cusp Sep | Length | meas | m_p_AV_Cusp_Sep | unilateral | 125195\17996-0\G-0394\T-35400 | SRT, T-35400, Aortic Valve | LN, 17996-0, Aortic Valve Cusp Separation |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Max PG | Maximum PG | meas | m_p_AV_Vmax_MaxPG | unilateral | 125195\20247-3\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Max PG (full) | none | calc | c_p_AV_MaxPG_FB | unilateral | 125195\125217\20247-3\T-35400 | SRT, T-35400, Aortic Valve | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury | [5 empty] | DCM, 125217, Full Bernoulli
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Mean PG | Mean PG | meas | m_p_AV_VTI_CardPGmean | unilateral | 125195\20256-4\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Mean PG (full) | none | calc | c_p_AV_MeanPG_FB | unilateral | 125195\125217\20256-4\T-35400 | SRT, T-35400, Aortic Valve | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury | [5 empty] | DCM, 125217, Full Bernoulli
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV R-R (DOP) | Time | meas | m_p_R_R_INT_AV | unilateral | 125195\8867-4\T-35400 | SRT, T-35400, Aortic Valve | LN, 8867-4, Heart Rate |  | UCUM, sec, Seconds
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV R-R (MM) | Time | meas | m_p_R_R_AV_MM | unilateral | 125195\8867-4\G-0394\T-35400 | SRT, T-35400, Aortic Valve | LN, 8867-4, Heart Rate |  | UCUM, sec, Seconds |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Vmax | Unsigned Maximum Velocity | meas | m_p_AV_Vmax_Vmax | unilateral | 125195\20351-3\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV Vmean | Unsigned Time Average Peak Vel | meas | m_p_AV_VTI_CardTAPV | unilateral | 125195\11692-1\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 11692-1, Time Averaged Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV VTI | Velocity Time Integral | meas | m_p_AV_VTI_VTI | unilateral | 125195\20354-7\R-42047\T-35400 | SRT, T-35400, Aortic Valve | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AVA (Vmax) | none | calc | c_p_AVA_Vmax | unilateral | 125195\125214\G-038E\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | [5 empty] | DCM, 125214, Continuity Equation by Peak Velocity
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AVA (VTI) | none | calc | c_p_AVA_VTI | unilateral | 125195\125215\G-038E\T-35400 | SRT, T-35400, Aortic Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | [5 empty] | DCM, 125215, Continuity Equation by Velocity Time Integral
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV-HR (DOP) | Heart Rate | meas | m_p_HR_AV | unilateral | 125195\8867-4\T-35400 | SRT, T-35400, Aortic Valve | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, Beats Per Minute
Pediatric Cardiac Ultrasound Report | Pediatric Echo | AV-HR (MM) | Heart Rate | meas | m_p_HR_AV_MM | unilateral | 125195\8867-4\G-0394\T-35400 | SRT, T-35400, Aortic Valve | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, Beats Per Minute |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | B-C Slope | MMode Slope | meas | m_p_BC_Slope_MMSlope | unilateral | 125195\59126-3\G-0394\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 59126-3, B-C Slope |  | UCUM, cm/s, Centimeter Per Second |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CI (2D-Cubed) | none | calc | c_p_CI_Cubed | unilateral | 125195\125206\F-32110\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CI (2D-Teich) | none | calc | c_p_CI_Teich | unilateral | 125195\125209\F-32110\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CI (A/L) | none | calc | c_p_CI_SPE | unilateral | 125195\125226\F-32110\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CI (A2C) | none | calc | c_p_CI_MOD_A2C | unilateral | 125195\125208\F-32110\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CI (A2C-A/L) | none | calc | c_p_CI_A2C_AL | unilateral | 125195\125226\F-32110\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CI (A4C) | none | calc | c_p_CI_MOD_A4C | unilateral | 125195\125208\F-32110\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CI (A4C-A/L) | none | calc | c_p_CI_A4C_AL | unilateral | 125195\125226\F-32110\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CI (BP) | none | calc | c_p_CI_MOD_BP | unilateral | 125195\125207\F-32110\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CI (MM-Cubed) | none | calc | c_p_CI_Cubed_MM | unilateral | 125195\125206\F-32110\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CI (MM-Teich) | none | calc | c_p_CI_Teich_MM | unilateral | 125195\125209\F-32110\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32110, Cardiac Index |  | UCUM, l/min/m2, l/min/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (2D-Cubed) | none | calc | c_p_CO_Cubed | unilateral | 125195\125206\F-32100\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (2D-Teich) | none | calc | c_p_CO_Teich | unilateral | 125195\125209\F-32100\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (A/L) | none | calc | c_p_CO_SPE | unilateral | 125195\125226\F-32100\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (A2C) | none | calc | c_p_CO_MOD_A2C | unilateral | 125195\125208\F-32100\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (A2C-A/L) | none | calc | c_p_CO_A2C_AL | unilateral | 125195\125226\F-32100\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (A4C) | none | calc | c_p_CO_MOD_A4C | unilateral | 125195\125208\F-32100\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (A4C-A/L) | none | calc | c_p_CO_A4C_AL | unilateral | 125195\125226\F-32100\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (BP) | none | calc | c_p_CO_MOD_BP | unilateral | 125195\125207\F-32100\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (LVOT) | none | calc | c_p_CO_LVOT | unilateral | 125195\F-32100\T-32600\T-32650 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute | SRT, T-32650, Left Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (MM-Cubed) | none | calc | c_p_CO_Cubed_MM | unilateral | 125195\125206\F-32100\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (MM-Teich) | none | calc | c_p_CO_Teich_MM | unilateral | 125195\125209\F-32100\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (MV) | none | calc | c_p_CO_MV | unilateral | 125195\F-32100\T-35300 | SRT, T-35300, Mitral Valve | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (PV) | none | calc | c_p_CO_PV | unilateral | 125195\F-32100\T-35200 | SRT, T-35200, Pulmonic Valve | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (RVOT) | none | calc | c_p_CO_RVOT | unilateral | 125195\F-32100\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute | SRT, T-32550, Right Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | CO (TV) | none | calc | c_p_CO_TV | unilateral | 125195\F-32100\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, F-32100, Cardiac Output |  | UCUM, l/min, Litre Per Minute
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Coarctation Diam | Length | meas | m_p_Coarc_Diam | unilateral | 125195\D4-32014\G-03A2\M-02550 | SRT, D4-32014, Coarctation of aorta | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Cx Diam | Length | meas | m_p_LCx_Diam | unilateral | 125195\G-03A2\M-02550\T-43120 | SRT, T-43120, Circumflex Coronary Artery | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Desc Ao Diam | Length | meas | m_p_DescAorta_Diam | unilateral | 125195\G-03A2\M-02550\T-42070 | SRT, T-42070, Thoracic aorta | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Desc Ao Max PG | Maximum PG | meas | m_p_DSC_Ao_MaxPG_MaxPG | unilateral | 125195\20247-3\T-42070 | SRT, T-42070, Thoracic aorta | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Desc Ao Mean PG | Mean PG | meas | m_p_DSC_Ao_PG | unilateral | 125195\20256-4\T-42070 | SRT, T-42070, Thoracic aorta | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Desc Ao Vmax | Unsigned Maximum Velocity | meas | m_p_DSC_Ao_MaxPG_Vmax | unilateral | 125195\20351-3\T-42070 | SRT, T-42070, Thoracic aorta | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Dsc Ao Max PG (full) | none | calc | c_p_DscAo_MaxPG_FB | unilateral | 125195\125217\20247-3\T-42070 | SRT, T-42070, Thoracic aorta | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury | [5 empty] | DCM, 125217, Full Bernoulli
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Dsc Ao Mean PG (full) | none | calc | c_p_DscAo_MeanPG_FB | unilateral | 125195\125217\20256-4\T-42070 | SRT, T-42070, Thoracic aorta | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury | [5 empty] | DCM, 125217, Full Bernoulli
Pediatric Cardiac Ultrasound Report | Pediatric Echo | E/E` Lateral | none | calc | c_p_E_LatE | unilateral | 125195\G-037B\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037B, Ratio of MV Peak Velocity to LV Peak Tissue Velocity E-Wave |  | UCUM, 1, no units | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | E/E` Medial | none | calc | c_p_E_MedE | unilateral | 125195\G-037B\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-037B, Ratio of MV Peak Velocity to LV Peak Tissue Velocity E-Wave |  | UCUM, 1, no units | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | E`/A` Lateral | none | calc | c_p_Lat_E_A | unilateral | 125195\59129-7\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59129-7, Left Ventricle E to A Tissue Velocity Ratio |  | UCUM, 1, no units | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | E`/A` Medial | none | calc | c_p_Med_E_A | unilateral | 125195\59129-7\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59129-7, Left Ventricle E to A Tissue Velocity Ratio |  | UCUM, 1, no units | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EDV (2D-Cubed) | none | calc | c_p_EDV_Cubed | unilateral | 125195\125206\F-32011\G-03A2\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EDV (2D-Teich) | none | calc | c_p_EDV_Teich | unilateral | 125195\125209\F-32011\G-03A2\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EDV (A/L) | none | calc | c_p_EDV_SPE | unilateral | 125195\125226\18026-5\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18026-5, Left Ventricular End Diastolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EDV (A2C) | MOD Volume | meas | m_p_EDV_MOD_A2C | unilateral | 125195\125208\F-32011\G-03A2\G-A19B\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EDV (A2C-A/L) | none | calc | c_p_EDV_A2C_AL | unilateral | 125195\125226\F-32011\G-03A2\G-A19B\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EDV (A4C) | MOD Volume | meas | m_p_EDV_MOD_A4C | unilateral | 125195\125208\F-32011\G-03A2\G-A19C\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EDV (A4C-A/L) | none | calc | c_p_EDV_A4C_AL | unilateral | 125195\125226\F-32011\G-03A2\G-A19C\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EDV (BP) | none | calc | c_p_EDV_MOD_BP | unilateral | 125195\125207\F-32011\G-03A2\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole |  | DCM, 125207, Method of Disks, Biplane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EDV (MM-Cubed) | none | calc | c_p_EDV_Cubed_MM | unilateral | 125195\125206\F-32011\G-0394\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  | UCUM, ml, Milliliter |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EDV (MM-Teich) | none | calc | c_p_EDV_Teich_MM | unilateral | 125195\125209\F-32011\G-0394\G-D705\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  | UCUM, ml, Milliliter |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EF (2D-Cubed) | none | calc | c_p_EF_Cubed | unilateral | 125195\125206\F-32070\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32070, Cardiac ejection fraction |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EF (2D-Teich) | none | calc | c_p_EF_Teich | unilateral | 125195\125209\F-32070\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32070, Cardiac ejection fraction |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EF (A/L) | none | calc | c_p_EF_SPE | unilateral | 125195\125226\F-32070\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32070, Cardiac ejection fraction |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EF (A2C) | none | calc | c_p_EF_MOD_A2C | unilateral | 125195\125208\F-32070\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32070, Cardiac ejection fraction |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EF (A2C-A/L) | none | calc | c_p_EF_A2C_AL | unilateral | 125195\125226\F-32070\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32070, Cardiac ejection fraction |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EF (A4C) | none | calc | c_p_EF_MOD_A4C | unilateral | 125195\125208\F-32070\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32070, Cardiac ejection fraction |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EF (A4C-A/L) | none | calc | c_p_EF_A4C_AL | unilateral | 125195\125226\F-32070\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32070, Cardiac ejection fraction |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EF (BP) | none | calc | c_p_EF_MOD_BP | unilateral | 125195\125207\F-32070\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32070, Cardiac ejection fraction |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EF (MM-Cubed) | none | calc | c_p_EF_Cubed_MM | unilateral | 125195\125206\F-32070\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32070, Cardiac ejection fraction |  | UCUM, %, Percent |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | EF (MM-Teich) | none | calc | c_p_EF_Teich_MM | unilateral | 125195\125209\F-32070\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32070, Cardiac ejection fraction |  | UCUM, %, Percent |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ESV (2D-Cubed) | none | calc | c_p_ESV_Cubed | unilateral | 125195\125206\G-03A2\G-D705\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  | SRT, R-FAB5B, End Systole |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ESV (2D-Teich) | none | calc | c_p_ESV_Teich | unilateral | 125195\125209\G-03A2\G-D705\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  | SRT, R-FAB5B, End Systole |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ESV (A/L) | none | calc | c_p_ESV_SPE | unilateral | 125195\125226\18148-7\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18148-7, Left Ventricular End Systolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ESV (A2C) | MOD Volume | meas | m_p_ESV_MOD_A2C | unilateral | 125195\125208\G-03A2\G-A19B\G-D705\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ESV (A2C-A/L) | none | calc | c_p_ESV_A2C_AL | unilateral | 125195\125226\G-03A2\G-A19B\G-D705\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ESV (A4C) | MOD Volume | meas | m_p_ESV_MOD_A4C | unilateral | 125195\125208\G-03A2\G-A19C\G-D705\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ESV (A4C-A/L) | none | calc | c_p_ESV_A4C_AL | unilateral | 125195\125226\G-03A2\G-A19C\G-D705\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ESV (BP) | none | calc | c_p_ESV_MOD_BP | unilateral | 125195\125207\G-03A2\G-D705\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  | SRT, R-FAB5B, End Systole |  | DCM, 125207, Method of Disks, Biplane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ESV (MM-Cubed) | none | calc | c_p_ESV_Cubed_MM | unilateral | 125195\125206\G-0394\G-D705\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  | UCUM, ml, Milliliter |  | SRT, G-0394, M mode |  | SRT, R-FAB5B, End Systole |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | ESV (MM-Teich) | none | calc | c_p_ESV_Teich_MM | unilateral | 125195\125209\G-0394\G-D705\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-D705, Volume |  | UCUM, ml, Milliliter |  | SRT, G-0394, M mode |  | SRT, R-FAB5B, End Systole |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Hepatic A Dur | Time | meas | m_p_HepARevsDur | unilateral | 125195\59105-7\T-48720 | SRT, T-48720, Hepatic Vein | LN, 59105-7, A-Wave Duration |  | UCUM, sec, Seconds
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Hepatic A Vel | Unsigned Maximum Velocity | meas | m_p_HepARevsVel_Vel | unilateral | 125195\59079-4\T-48720 | SRT, T-48720, Hepatic Vein | LN, 59079-4, Peak Reversal Velocity during Atrial Contraction |  | UCUM, cm/s, Centimeter Per Second
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Hepatic Dias Vel | Unsigned Maximum Velocity | meas | m_p_HepDiasVel_Vel | unilateral | 125195\20351-3\F-32011\T-48720 | SRT, T-48720, Hepatic Vein | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Hepatic S/D | none | calc | c_p_HepSD | unilateral | 125195\12144-2\T-48720 | SRT, T-48720, Hepatic Vein | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Hepatic Sys Vel | Unsigned Maximum Velocity | meas | m_p_HepSysVel_Vel | unilateral | 125195\20351-3\R-FAB5B\T-48720 | SRT, T-48720, Hepatic Vein | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | HR LV (2D) | Heart Rate | meas | m_p_HR_LV | unilateral | 125195\8867-4\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, Beats Per Minute |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | HR LV (MM) | Heart Rate | meas | m_p_HR_LV_MM | unilateral | 125195\8867-4\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, Beats Per Minute |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVC A Dur | Time | meas | m_p_IVC_ARevsDur | unilateral | 125195\59105-7\T-48710 | SRT, T-48710, Inferior vena cava | LN, 59105-7, A-Wave Duration |  | UCUM, sec, Seconds
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVC A Vel | Unsigned Maximum Velocity | meas | m_p_IVC_ARevsVel_Vel | unilateral | 125195\59079-4\T-48710 | SRT, T-48710, Inferior vena cava | LN, 59079-4, Peak Reversal Velocity during Atrial Contraction |  | UCUM, cm/s, Centimeter Per Second
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVC Diam | Length | meas | m_p_IVC_Diam | unilateral | 125195\G-03A2\M-02550\T-48710 | SRT, T-48710, Inferior vena cava | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVC Dias Vel | Unsigned Maximum Velocity | meas | m_p_IVC_DiasVel_Vel | unilateral | 125195\20351-3\F-32011\T-48710 | SRT, T-48710, Inferior vena cava | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVC S/D | none | calc | c_p_IVC_SD | unilateral | 125195\12144-2\T-48710 | SRT, T-48710, Inferior vena cava | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVC Sys Vel | Unsigned Maximum Velocity | meas | m_p_IVC_SysVel_Vel | unilateral | 125195\20351-3\R-FAB5B\T-48710 | SRT, T-48710, Inferior vena cava | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVCT | Time | meas | m_p_IVCT | unilateral | 125195\59084-4\T-32600 | SRT, T-32600, Left Ventricle | LN, 59084-4, Isovolumic Contraction Time |  | UCUM, sec, Seconds
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVRT | Time | meas | m_p_IVRT | unilateral | 125195\59083-6\T-32600 | SRT, T-32600, Left Ventricle | LN, 59083-6, Isovolumic Relaxation Time |  | UCUM, sec, Seconds
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVS % (2D) | none | calc | c_p_IVS | unilateral | 125195\59092-7\G-03A2\T-32410 | SRT, T-32410, Interventricular septum | LN, 59092-7, % Thickening |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVS % (MM) | none | calc | c_p_IVS_MM | unilateral | 125195\59092-7\G-0394\T-32410 | SRT, T-32410, Interventricular septum | LN, 59092-7, % Thickening |  | UCUM, %, Percent |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVS/LVPW (2D) | none | calc | c_p_IVS_LVPW | unilateral | 125195\18155-2\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 18155-2, Interventricular Septum to Posterior Wall Thickness Ratio |  | UCUM, 1, no units |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVS/LVPW (MM) | none | calc | c_p_IVS_LVPW_MM | unilateral | 125195\18155-2\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18155-2, Interventricular Septum to Posterior Wall Thickness Ratio |  | UCUM, 1, no units |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVSd (2D) | Length | meas | m_p_IVSd_2D | unilateral | 125195\59089-3\F-32011\G-03A2\T-32410 | SRT, T-32410, Interventricular septum | LN, 59089-3, ROI Thickness by US |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVSd (MM) | Length | meas | m_p_IVSd_MM | unilateral | 125195\59089-3\F-32011\G-0394\T-32410 | SRT, T-32410, Interventricular septum | LN, 59089-3, ROI Thickness by US |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVSs (2D) | Length | meas | m_p_IVSs_2D | unilateral | 125195\59089-3\G-03A2\R-FAB5B\T-32410 | SRT, T-32410, Interventricular septum | LN, 59089-3, ROI Thickness by US |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | IVSs (MM) | Length | meas | m_p_IVSs_MM | unilateral | 125195\59089-3\G-0394\R-FAB5B\T-32410 | SRT, T-32410, Interventricular septum | LN, 59089-3, ROI Thickness by US |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA Dimen (2D) | Length | meas | m_p_LA_Dimen_2D | unilateral | 125195\29469-4\G-03A2\R-FAB5B\T-32300 | SRT, T-32300, Left Atrium | LN, 29469-4, Left Atrium Antero-posterior Systolic Dimension |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA Dimen (MM) | Length | meas | m_p_LA_Dimen_MM | unilateral | 125195\29469-4\G-0394\R-FAB5B\T-32300 | SRT, T-32300, Left Atrium | LN, 29469-4, Left Atrium Antero-posterior Systolic Dimension |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA ESV (BP) | none | calc | c_p_LA_ESV_BP | unilateral | 125195\125207\G-0383\G-03A2\T-32300 | SRT, T-32300, Left Atrium | SRT, G-0383, Left Atrium Systolic Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA ESV/BSA (A2C) | none | calc | c_p_LA_ESV_A2C_BSA | unilateral | 125195\125208\8277-6\C12205-03\G-03A2\G-A19B\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-03, Left Atrium Systolic Volume Index |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA ESV/BSA (A4C) | none | calc | c_p_LA_ESV_A4C_BSA | unilateral | 125195\125208\8277-6\C12205-03\G-03A2\G-A19C\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-03, Left Atrium Systolic Volume Index |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA ESV/BSA (BP) | none | calc | c_p_LA_ESV_BP_BSA | unilateral | 125195\125207\8277-6\C12205-03\G-03A2\T-32300 | SRT, T-32300, Left Atrium | 99PMSBLUS, C12205-03, Left Atrium Systolic Volume Index |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA/Ao (2D) | none | calc | c_p_LA_Ao_2D | unilateral | 125195\17985-3\G-03A2\T-32300 | SRT, T-32300, Left Atrium | LN, 17985-3, Left Atrium to Aortic Root Ratio |  | UCUM, 1, no units |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LA/Ao (MM) | none | calc | c_p_LA_Ao_MM | unilateral | 125195\17985-3\G-0394\T-32300 | SRT, T-32300, Left Atrium | LN, 17985-3, Left Atrium to Aortic Root Ratio |  | UCUM, 1, no units |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LAD Diam | Length | meas | m_p_LAD_Diam | unilateral | 125195\G-03A2\M-02550\T-43110 | SRT, T-43110, Anterior Descending Branch of Left Coronary Artery | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LAED Major-A4C | Length | meas | m_p_LAED_MAJOR_A4C | unilateral | 125195\F-32011\G-03A2\G-A193\G-A19C\T-32300 | SRT, T-32300, Left Atrium | SRT, G-A193, Major Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LAED Minor-A4C | Length | meas | m_p_LAED_MINOR_A4C | unilateral | 125195\F-32011\G-03A2\G-A194\G-A19C\T-32300 | SRT, T-32300, Left Atrium | SRT, G-A194, Minor Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LAES Major-A4C | Length | meas | m_p_LAES_MAJOR_A4C | unilateral | 125195\G-03A2\G-A193\G-A19C\R-FAB5B\T-32300 | SRT, T-32300, Left Atrium | SRT, G-A193, Major Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LAES Minor-A4C | Length | meas | m_p_LAES_MINOR_A4C | unilateral | 125195\G-03A2\G-A194\G-A19C\R-FAB5B\T-32300 | SRT, T-32300, Left Atrium | SRT, G-A194, Minor Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Lat A` Area | Velocity Time Integral | meas | m_p_AreaUnderLatA | unilateral | 125195\59125-5\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59125-5, LV VTI A wave US |  | UCUM, cm, Centimeter | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Lat A` Vel | Unsigned Maximum Velocity | meas | m_p_Lat_AVel | unilateral | 125195\59133-9\F-32030\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59133-9, Peak Tissue Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging |  | SRT, F-32030, Atrial Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Lat Accel Time | Time | meas | m_p_TDILatAccelTime | unilateral | 125195\20168-1\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Lat Decel Time | Time | meas | m_p_TDILatDecelTime | unilateral | 125195\20217-6\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 20217-6, Deceleration Time |  | UCUM, sec, Seconds | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Lat E` Area | Velocity Time Integral | meas | m_p_AreaUnderLatE | unilateral | 125195\59124-8\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59124-8, LV VTI E wave US |  | UCUM, cm, Centimeter | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Lat E` Vel | Unsigned Maximum Velocity | meas | m_p_Lat_EVel | unilateral | 125195\59133-9\G-0392\P5-B0128\R-40B1B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59133-9, Peak Tissue Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging |  | SRT, R-40B1B, Early Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Lat IVCT | Time | meas | m_p_TDI_Lat_IVCT | unilateral | 125195\59084-4\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59084-4, Isovolumic Contraction Time |  | UCUM, sec, Seconds | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Lat IVRT | Time | meas | m_p_TDI_Lat_IVRT | unilateral | 125195\59083-6\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59083-6, Isovolumic Relaxation Time |  | UCUM, sec, Seconds | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Lat S Vel | Unsigned Maximum Velocity | meas | m_p_Lat_SVel | unilateral | 125195\59133-9\G-0392\P5-B0128\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59133-9, Peak Tissue Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Late Dias Slope | MMode Slope | meas | m_p_LateDiasSlope_MMSlope | unilateral | 125195\20216-8\F-32011\G-0394\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20216-8, Deceleration Slope |  | UCUM, cm/s, Centimeter Per Second |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Left Main | Length | meas | m_p_L_Main_Diam | unilateral | 125195\G-03A2\M-02550\T-43107 | SRT, T-43107, Left Main Coronary Artery | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LL PulmV A Dur | Time | meas | m_p_LL_PulmARevsDur | unilateral | 125195\59105-7\T-48540\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 59105-7, A-Wave Duration |  | UCUM, sec, Seconds | SRT, T-48540, Left Inferior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LL PulmV A Vel | Unsigned Maximum Velocity | meas | m_p_LL_PulmARevsVel_Vel | unilateral | 125195\59079-4\T-48540\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 59079-4, Peak Reversal Velocity during Atrial Contraction |  | UCUM, cm/s, Centimeter Per Second | SRT, T-48540, Left Inferior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LL PulmV Diam | Length | meas | m_p_LL_PulmV_Diam | unilateral | 125195\G-03A2\M-02550\T-48540\T-48581 | SRT, T-48581, Pulmonary Vein | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter | SRT, T-48540, Left Inferior Pulmonary Vein | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LL PulmV Dias Vel | Unsigned Maximum Velocity | meas | m_p_LL_PulmDiasVel_Vel | unilateral | 125195\20351-3\F-32011\T-48540\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, T-48540, Left Inferior Pulmonary Vein |  |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LL PulmV S/D | none | calc | c_p_LL_PulmSD | unilateral | 125195\12144-2\T-48540\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | SRT, T-48540, Left Inferior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LL PulmV Sys Vel | Unsigned Maximum Velocity | meas | m_p_LL_PulmSysVel_Vel | unilateral | 125195\20351-3\R-FAB5B\T-48540\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, T-48540, Left Inferior Pulmonary Vein |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LPA Diam | Length | meas | m_p_LPA_Diam | unilateral | 125195\G-03A2\M-02550\T-44000\T-44400 | SRT, T-44000, Pulmonary artery | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter | SRT, T-44400, Left pulmonary artery | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LPA Max PG | Maximum PG | meas | m_p_LPA_Vmax_MaxPG | unilateral | 125195\20247-3\T-44000\T-44400 | SRT, T-44000, Pulmonary artery | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury | SRT, T-44400, Left pulmonary artery
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LPA Mean PG | Mean PG | meas | m_p_LPA_PG | unilateral | 125195\20256-4\T-44000\T-44400 | SRT, T-44000, Pulmonary artery | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury | SRT, T-44400, Left pulmonary artery
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LPA Vmax | Unsigned Maximum Velocity | meas | m_p_LPA_Vmax_Vmax | unilateral | 125195\20351-3\T-44000\T-44400 | SRT, T-44000, Pulmonary artery | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, T-44400, Left pulmonary artery
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LU PulmV A Dur | Time | meas | m_p_LU_PulmARevsDur | unilateral | 125195\59105-7\T-48530\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 59105-7, A-Wave Duration |  | UCUM, sec, Seconds | SRT, T-48530, Left Superior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LU PulmV A Vel | Unsigned Maximum Velocity | meas | m_p_LU_PulmARevsVel_Vel | unilateral | 125195\59079-4\T-48530\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 59079-4, Peak Reversal Velocity during Atrial Contraction |  | UCUM, cm/s, Centimeter Per Second | SRT, T-48530, Left Superior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LU PulmV Diam | Length | meas | m_p_LU_PulmV_Diam | unilateral | 125195\G-03A2\M-02550\T-48530\T-48581 | SRT, T-48581, Pulmonary Vein | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter | SRT, T-48530, Left Superior Pulmonary Vein | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LU PulmV Dias Vel | Unsigned Maximum Velocity | meas | m_p_LU_PulmDiasVel_Vel | unilateral | 125195\20351-3\F-32011\T-48530\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, T-48530, Left Superior Pulmonary Vein |  |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LU PulmV S/D | none | calc | c_p_LU_PulmSD | unilateral | 125195\12144-2\T-48530\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | SRT, T-48530, Left Superior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LU PulmV Sys Vel | Unsigned Maximum Velocity | meas | m_p_LU_PulmSysVel_Vel | unilateral | 125195\20351-3\R-FAB5B\T-48530\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, T-48530, Left Superior Pulmonary Vein |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV dP/dt | Doppler Dp/dt | meas | m_p_LV_dpdt | unilateral | 125195\18035-6\T-35300 | SRT, T-35300, Mitral Valve | LN, 18035-6, Mitral Regurgitation dP/dt derived from Mitral Reg velocity |  | UCUM, mm[Hg]/s, mmHg/s
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV ET (DOP) | Time | meas | m_p_LV_ET_Dop | unilateral | 125195\20222-6\T-32600 | SRT, T-32600, Left Ventricle | LN, 20222-6, Ejection Time |  | UCUM, sec, Seconds
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV ET (MM) | Time | meas | m_p_LV_Eject_Tm | unilateral | 125195\20222-6\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 20222-6, Ejection Time |  | UCUM, sec, Seconds |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV FS (2D) | none | calc | c_p_LVFS_2D | unilateral | 125195\59132-1\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 59132-1, Fractional Shortening |  | UCUM, %, Percent |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV FS (MM) | none | calc | c_p_LVFS_MM | unilateral | 125195\59132-1\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59132-1, Fractional Shortening |  | UCUM, %, Percent |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV Mass (A/L) | none | calc | c_p_LVMass_A_L | unilateral | 125195\125270\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | DCM, 125270, Left Ventricle Mass by Area Length |  | UCUM, g, Gram |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV Mass (Cubed) | none | calc | c_p_LV_Mass | unilateral | 125195\125206\18087-7\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 18087-7, Left Ventricle Mass |  | UCUM, g, Gram |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV Mass Index (A/L) | none | calc | c_p_LVMass_A_LBSA | unilateral | 125195\8277-6\C12203-01\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-01, Left Ventricle Mass Index |  | UCUM, g/m2, g/m2 |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV Mass Index(Cubed) | none | calc | c_p_LV_Mass_BSA | unilateral | 125195\125206\8277-6\C12203-01\G-0394\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, C12203-01, Left Ventricle Mass Index |  | UCUM, g/m2, g/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV MPI | none | calc | c_p_LV_MPI | unilateral | 125195\59099-2\T-32600 | SRT, T-32600, Left Ventricle | LN, 59099-2, Myocardial Performance Index (Tei) |  | UCUM, 1, no units
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV PEP | Time | meas | m_p_LV_PreEject_Tm | unilateral | 125195\59085-1\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59085-1, Pre-Ejection Period |  | UCUM, sec, Seconds |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV PEP/ET | none | calc | c_p_LV_PEP_ET | unilateral | 125195\59088-5\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59088-5, Pre-Ejection Period/Ejection Time Ratio |  | UCUM, 1, no units |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV PEPc | none | calc | c_p_LVPEPc | unilateral | 125195\59087-7\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59087-7, Heart Rate-Corrected Pre-Ejection Period |  | UCUM, ms, Millisecond |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV R-R (2D) | Time | meas | m_p_R_R_INT_LV | unilateral | 125195\8867-4\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 8867-4, Heart Rate |  | UCUM, sec, Seconds |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LV R-R (MM) | Time | meas | m_p_MM_DIM_R_R_INT | unilateral | 125195\8867-4\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 8867-4, Heart Rate |  | UCUM, sec, Seconds |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVAd (A/L) | Area | meas | m_p_LVAd_Apical | unilateral | 125195\125226\F-32011\G-03A2\G-A166\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A166, Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVAd (A2C-A/L) | Area | meas | m_p_LVAD_A2C | unilateral | 125195\125208\F-32011\G-0375\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0375, Left Ventricular Diastolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVAd (A4C-A/L) | Area | meas | m_p_LVAD_A4C | unilateral | 125195\125208\F-32011\G-0375\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0375, Left Ventricular Diastolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVAd Sax Endo Area | Area | meas | m_p_LV_AdSaxEndo_Area | unilateral | 125195\59094-3\F-32011\G-039B\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 59094-3, Endocardial Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-039B, Parasternal short axis at the Papillary Muscle level | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVAd Sax Endo Circ | Circumference | meas | m_p_LV_AdSaxEndo_Circ | unilateral | 125195\59094-3-1\F-32011\G-039B\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, 59094-3-1, Endocardial Circumference |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-039B, Parasternal short axis at the Papillary Muscle level | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVAd Sax Epi Area | Area | meas | m_p_LV_AdSaxEpi_Area | unilateral | 125195\59093-5\F-32011\G-039B\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 59093-5, Epicardial Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-039B, Parasternal short axis at the Papillary Muscle level | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVAd Sax Epi Circ | Circumference | meas | m_p_LV_AdSaxEpi_Circ | unilateral | 125195\59093-5-1\F-32011\G-039B\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | 99PMSBLUS, 59093-5-1, Epicardial Circumference |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-039B, Parasternal short axis at the Papillary Muscle level | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVAs (A/L) | Area | meas | m_p_LVAs_Apical | unilateral | 125195\125226\G-03A2\G-A166\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A166, Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode |  | SRT, R-FAB5B, End Systole |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVAs (A2C-A/L) | Area | meas | m_p_LVAS_A2C | unilateral | 125195\125208\G-0374\G-03A2\G-A19B\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0374, Left Ventricular Systolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVAs (A4C-A/L) | Area | meas | m_p_LVAS_A4C | unilateral | 125195\125208\G-0374\G-03A2\G-A19C\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-0374, Left Ventricular Systolic Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVDP (AI) | none | calc | c_p_LVDP_AI | unilateral | 125195\C12212-02\F-32011\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12212-02, Left Ventricle Diastolic Pressure with Aortic Insufficiency |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVED Major-A4C | Length | meas | m_p_LVED_MAJOR_A4C | unilateral | 125195\F-32011\G-03A2\G-A193\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVED Minor-A4C | Length | meas | m_p_LVED_MINOR_A4C | unilateral | 125195\F-32011\G-03A2\G-A194\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A194, Minor Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVED SAX CH | Length | meas | m_p_LVED_SAX_CH | unilateral | 125195\121206\F-32011\G-0399\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | DCM, 121206, Distance |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-0399, Parasternal short axis at the level of the mitral chords | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVED SAX PM | Length | meas | m_p_LVED_SAX_PM | unilateral | 125195\121206\F-32011\G-039B\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | DCM, 121206, Distance |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-039B, Parasternal short axis at the Papillary Muscle level | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVES Major-A4C | Length | meas | m_p_LVES_MAJOR_A4C | unilateral | 125195\G-03A2\G-A193\G-A19C\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVES Minor-A4C | Length | meas | m_p_LVES_MINOR_A4C | unilateral | 125195\G-03A2\G-A194\G-A19C\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A194, Minor Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVES SAX CH | Length | meas | m_p_LVES_SAX_CH | unilateral | 125195\121206\G-0399\G-03A2\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | DCM, 121206, Distance |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-0399, Parasternal short axis at the level of the mitral chords | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVES SAX PM | Length | meas | m_p_LVES_SAX_PM | unilateral | 125195\121206\G-039B\G-03A2\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | DCM, 121206, Distance |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-039B, Parasternal short axis at the Papillary Muscle level | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVETc | none | calc | c_p_LV_ETc | unilateral | 125195\59086-9\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59086-9, Heart Rate-Corrected Ejection Time |  | UCUM, sec, Seconds |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVIDd (2D) | Length | meas | m_p_LVIDd_2D | unilateral | 125195\59090-1\F-32011\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 59090-1, ROI Internal Dimension by US |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVIDd (MM) | Length | meas | m_p_LVIDd_MM | unilateral | 125195\59090-1\F-32011\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59090-1, ROI Internal Dimension by US |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVIDs (2D) | Length | meas | m_p_LVIDs_2D | unilateral | 125195\59090-1\G-03A2\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59090-1, ROI Internal Dimension by US |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVIDs (MM) | Length | meas | m_p_LVIDs_MM | unilateral | 125195\59090-1\G-0394\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59090-1, ROI Internal Dimension by US |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVLd (A/L) | Length | meas | m_p_LVLd_Apical | unilateral | 125195\125226\F-32011\G-03A2\G-A193\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVLd (A2C-A/L) | Length | meas | m_p_LVLD_A2C | unilateral | 125195\125208\18077-8\F-32011\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | LN, 18077-8, Left Ventricle diastolic major axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVLd (A4C-A/L) | Length | meas | m_p_LVLD_A4C | unilateral | 125195\125208\18077-8\F-32011\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | LN, 18077-8, Left Ventricle diastolic major axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVLd Apical | Length | meas | m_p_LV_LdApical_Distance | unilateral | 125195\F-32011\G-03A2\G-A193\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVLs (A/L) | Length | meas | m_p_LVLs_Apical | unilateral | 125195\125226\G-03A2\G-A193\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | SRT, G-A193, Major Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode |  | SRT, R-FAB5B, End Systole |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVLs (A2C-A/L) | Length | meas | m_p_LVLS_A2C | unilateral | 125195\125208\18076-0\G-03A2\G-A19B\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 18076-0, Left Ventricle systolic major axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVLs (A4C-A/L) | Length | meas | m_p_LVLS_A4C | unilateral | 125195\125208\18076-0\G-03A2\G-A19C\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 18076-0, Left Ventricle systolic major axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVOT Accel Slope | Doppler Slope | meas | m_p_LVOT_AccelSlope_DopSlope | unilateral | 125195\20167-3\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20167-3, Acceleration Slope |  | UCUM, cm/s2, Centimeter Per Second Square | SRT, T-32650, Left Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVOT Accel Time | Time | meas | m_p_LVOT_AccelSlope_DopTime | unilateral | 125195\20168-1\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | SRT, T-32650, Left Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVOT Area | Area | meas | m_p_LVOT_Diam_Area | unilateral | 125195\G-038E\T-32600\T-32650 | SRT, T-32600, Left Ventricle | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | SRT, T-32650, Left Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVOT Area | none | calc | c_p_LVOT_Area | unilateral | 125195\G-038E\G-03A2\T-32600\T-32650 | SRT, T-32600, Left Ventricle | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | SRT, T-32650, Left Ventricle Outflow Tract | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVOT Diam | Length | meas | m_p_LVOT_Diam | unilateral | 125195\G-038F\T-32600\T-32650 | SRT, T-32600, Left Ventricle | SRT, G-038F, Cardiovascular Orifice Diameter |  | UCUM, cm, Centimeter | SRT, T-32650, Left Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVOT Max PG | Maximum PG | meas | m_p_LVOT_Vmax_MaxPG | unilateral | 125195\20247-3\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury | SRT, T-32650, Left Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVOT Mean PG | Mean PG | meas | m_p_LVOT_VTI_CardPGmean | unilateral | 125195\20256-4\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury | SRT, T-32650, Left Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVOT Vmax | Unsigned Maximum Velocity | meas | m_p_LVOT_Vmax_Vmax | unilateral | 125195\20351-3\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, T-32650, Left Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVOT Vmean | Unsigned Time Average Peak Vel | meas | m_p_LVOT_VTI_CardTAPV | unilateral | 125195\11692-1\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 11692-1, Time Averaged Peak Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, T-32650, Left Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVOT VTI | Velocity Time Integral | meas | m_p_LVOT_VTI_VTI | unilateral | 125195\20354-7\T-32600\T-32650 | SRT, T-32600, Left Ventricle | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | SRT, T-32650, Left Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVPW % (2D) | none | calc | c_p_LVPW | unilateral | 125195\59092-7\C12243-01\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | LN, 59092-7, % Thickening |  | UCUM, %, Percent | 99PMSBLUS, C12243-01, Left Ventricle Posterior Wall | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVPW % (MM) | none | calc | c_p_LVPW_MM | unilateral | 125195\59092-7\C12243-01\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59092-7, % Thickening |  | UCUM, %, Percent | 99PMSBLUS, C12243-01, Left Ventricle Posterior Wall | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVPWd (2D) | Length | meas | m_p_LVPWd_2D | unilateral | 125195\59089-3\F-32011\G-03A2\R-42175\T-32600 | SRT, T-32600, Left Ventricle | LN, 59089-3, ROI Thickness by US |  | UCUM, cm, Centimeter | SRT, R-42175, Posterior Wall | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVPWd (MM) | Length | meas | m_p_LVPWd_MM | unilateral | 125195\59089-3\F-32011\G-0394\R-42175\T-32600 | SRT, T-32600, Left Ventricle | LN, 59089-3, ROI Thickness by US |  | UCUM, cm, Centimeter | SRT, R-42175, Posterior Wall | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVPWs (2D) | Length | meas | m_p_LVPWs_2D | unilateral | 125195\59089-3\G-03A2\R-42175\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59089-3, ROI Thickness by US |  | UCUM, cm, Centimeter | SRT, R-42175, Posterior Wall | SRT, G-03A2, 2D mode |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVPWs (MM) | Length | meas | m_p_LVPWs_MM | unilateral | 125195\59089-3\G-0394\R-42175\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59089-3, ROI Thickness by US |  | UCUM, cm, Centimeter | SRT, R-42175, Posterior Wall | SRT, G-0394, M mode |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | LVSP (AS) | none | calc | c_p_LVSP_AS | unilateral | 125195\C12212-01\R-FAB5B\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12212-01, Left Ventricle Systolic Pressure with Aortic Stenosis |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Mean VCF | none | calc | c_p_Mean_VCF | unilateral | 125195\59117-2\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59117-2, Mean Velocity of Circumferential Fiber Shortening (Mean VcFv) |  | 99PMSBLUS, circ/s, circ/s |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Mean VCFc | none | calc | c_p_Mean_VCFc | unilateral | 125195\59118-0\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59118-0, HR-Corrected Mean Velocity of Circumferential Fiber Shortening |  | 99PMSBLUS, circ/s, circ/s |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Med A` Area | Velocity Time Integral | meas | m_p_AreaUnderMedA | unilateral | 125195\59125-5\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59125-5, LV VTI A wave US |  | UCUM, cm, Centimeter | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Med A` Vel | Unsigned Maximum Velocity | meas | m_p_Med_AVel | unilateral | 125195\59133-9\F-32030\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59133-9, Peak Tissue Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging |  | SRT, F-32030, Atrial Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Med Accel Time | Time | meas | m_p_TDIMedAccelTime | unilateral | 125195\20168-1\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Med Decel Time | Time | meas | m_p_TDIMedDecelTime | unilateral | 125195\20217-6\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 20217-6, Deceleration Time |  | UCUM, sec, Seconds | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Med E` Area | Velocity Time Integral | meas | m_p_AreaUnderMedE | unilateral | 125195\59124-8\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59124-8, LV VTI E wave US |  | UCUM, cm, Centimeter | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Med E` Vel | Unsigned Maximum Velocity | meas | m_p_Med_EVel | unilateral | 125195\59133-9\G-0391\P5-B0128\R-40B1B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59133-9, Peak Tissue Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging |  | SRT, R-40B1B, Early Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Med IVCT | Time | meas | m_p_TDI_Med_IVCT | unilateral | 125195\59084-4\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59084-4, Isovolumic Contraction Time |  | UCUM, sec, Seconds | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Med IVRT | Time | meas | m_p_TDI_Med_IVRT | unilateral | 125195\59083-6\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59083-6, Isovolumic Relaxation Time |  | UCUM, sec, Seconds | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Med S Vel | Unsigned Maximum Velocity | meas | m_p_Med_SVel | unilateral | 125195\59133-9\G-0391\P5-B0128\R-FAB5B\T-32600 | SRT, T-32600, Left Ventricle | LN, 59133-9, Peak Tissue Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MPA Area | none | calc | c_p_MPA_Area | unilateral | 125195\G-038E\G-03A2\T-44100 | SRT, T-44100, Pulmonary Trunk | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MPA Diam | Length | meas | m_p_MPA_Diam | unilateral | 125195\G-03A2\M-02550\T-44000 | SRT, T-44000, Pulmonary artery | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MPA Max PG | Maximum PG | meas | m_p_MPA_Vmax_MaxPG | unilateral | 125195\20247-3\T-44100 | SRT, T-44100, Pulmonary Trunk | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MPA Mean PG | Mean PG | meas | m_p_MPA_PG | unilateral | 125195\20256-4\T-44100 | SRT, T-44100, Pulmonary Trunk | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MPA Vmax | Unsigned Maximum Velocity | meas | m_p_MPA_Vmax_Vmax | unilateral | 125195\20351-3\T-44100 | SRT, T-44100, Pulmonary Trunk | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MR Alias Vel | Unsigned Maximum Velocity | meas | m_p_MR_AliasVel | unilateral | 125195\59130-5\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 59130-5, Alias velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MR ERO | none | calc | c_p_MR_ERO | unilateral | 125195\125216\G-0367\G-038E\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MR Flow Rate | none | calc | c_p_MR_FlowRate | unilateral | 125195\34141-2\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 34141-2, Peak Instantaneous Flow Rate |  | UCUM, ml/sec, ml/sec |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MR Fraction | none | calc | c_p_MR_Frac | unilateral | 125195\G-0367\G-0390\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-0390, Regurgitant Fraction |  | UCUM, %, Percent |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MR Max PG | Maximum PG | meas | m_p_MR_Vmax_MaxPG | unilateral | 125195\20247-3\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MR Mean PG | Mean PG | meas | m_p_MR_VTI_CardPGmean | unilateral | 125195\20256-4\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MR Radius | Length | meas | m_p_MR_Radius | unilateral | 125195\59102-4\G-0367\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | LN, 59102-4, Flow Radius |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MR Vmax | Unsigned Maximum Velocity | meas | m_p_MR_Vmax_Vmax | unilateral | 125195\20351-3\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MR Vmean | Unsigned Time Average Peak Vel | meas | m_p_MR_VTI_CardTAPV | unilateral | 125195\11692-1\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 11692-1, Time Averaged Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MR Volume | none | calc | c_p_MR_Vol | unilateral | 125195\125216\33878-0\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 33878-0, Volume Flow |  | UCUM, ml, Milliliter |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MR VTI | Velocity Time Integral | meas | m_p_MR_VTI_VTI | unilateral | 125195\20354-7\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV A Dur | Time | meas | m_p_MV_A_Dur | unilateral | 125195\59105-7\T-35300 | SRT, T-35300, Mitral Valve | LN, 59105-7, A-Wave Duration |  | UCUM, sec, Seconds
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV A-C Interval | Time | meas | m_p_MV_AC_Int | unilateral | 125195\59103-2\G-0394\T-35300 | SRT, T-35300, Mitral Valve | LN, 59103-2, A-C Interval |  | UCUM, sec, Seconds |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Acc Slope | Doppler Slope | meas | m_p_MV_AS_DopSlope | unilateral | 125195\20167-3\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20167-3, Acceleration Slope |  | UCUM, cm/s2, Centimeter Per Second Square |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Acc Time | Time | meas | m_p_MV_AS_DopTime | unilateral | 125195\20168-1\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Alias Vel | Unsigned Maximum Velocity | meas | m_p_MV_AliasVel | unilateral | 125195\59130-5\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 59130-5, Alias velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Annul Area | none | calc | c_p_MV_Annul_Area | unilateral | 125195\G-038E\G-03A2\T-3500E\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | SRT, T-3500E, Cardiac valve annulus | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Annul Diam | Length | meas | m_p_MV_Annul_Diam | unilateral | 125195\G-03A2\M-02550\T-3500E\T-35300 | SRT, T-35300, Mitral Valve | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter | SRT, T-3500E, Cardiac valve annulus | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Area | none | calc | c_p_MV_Area | unilateral | 125195\G-03A2\G-A166\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-A166, Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Area (Ellipse) | none | calc | c_p_MV_Area_Ellps | unilateral | 125195\125211\G-038E\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125211, Biplane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Area (Planim) | Area | meas | m_p_MV_Area | unilateral | 125195\125220\G-038E\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125220, Planimetry
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Closure to Opening | Time | meas | m_p_MV_Cl_to_Op | unilateral | 125195\59082-8\T-35300 | SRT, T-35300, Mitral Valve | LN, 59082-8, Closure to Opening Time |  | UCUM, sec, Seconds
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV D-E Exc Dist | Length | meas | m_p_MV_DE_Exc_Dist | unilateral | 125195\59091-9\G-0394\T-35300 | SRT, T-35300, Mitral Valve | LN, 59091-9, D-E Excursion |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV D-E Slope | MMode Slope | meas | m_p_MV_DE_Slope_MMSlope | unilateral | 125195\59127-1\G-0394\T-35300 | SRT, T-35300, Mitral Valve | LN, 59127-1, D-E Slope |  | UCUM, cm/s, Centimeter Per Second |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Decel Slope | Doppler Slope | meas | m_p_MV_DS_DopSlope | unilateral | 125195\20216-8\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20216-8, Deceleration Slope |  | UCUM, cm/s2, Centimeter Per Second Square |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Decel Time | Time | meas | m_p_MV_DS_DopTime | unilateral | 125195\20217-6\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20217-6, Deceleration Time |  | UCUM, sec, Seconds |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV DFP | Time | meas | m_p_MV_DFP | unilateral | 125195\R-0032C\T-35300 | SRT, T-35300, Mitral Valve | SRT, R-0032C, Mitral Diastolic Filling Period (DFPm) |  | UCUM, sec, Seconds
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Diam | Length | meas | m_p_MV_Diam | unilateral | 125195\G-038F\G-03A2\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038F, Cardiovascular Orifice Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV E/A | none | calc | c_p_MV_EA | unilateral | 125195\59104-0\P5-B0128\T-35300 | SRT, T-35300, Mitral Valve | LN, 59104-0, Peak E wave/Peak A wave by US |  | UCUM, 1, no units |  | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV E-E Sep | Length | meas | m_p_MV_EE_Separ | unilateral | 125195\59098-4\G-0394\T-35300 | SRT, T-35300, Mitral Valve | LN, 59098-4, Mitral Valve E-septal Separation |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV E-F Slope | MMode Slope | meas | m_p_MV_EF_Slope_MMSlope | unilateral | 125195\59128-9\G-0394\T-35300 | SRT, T-35300, Mitral Valve | LN, 59128-9, E-F Slope |  | UCUM, cm/s, Centimeter Per Second |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV EPSS | Length | meas | m_p_MV_EPSS | unilateral | 125195\18036-4\G-0394\T-35300 | SRT, T-35300, Mitral Valve | LN, 18036-4, Mitral Valve EPSS, E wave |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Major | Length | meas | m_p_MV_Major | unilateral | 125195\G-03A2\G-A193\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-A193, Major Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Max PG | Maximum PG | meas | m_p_MV_Vmax_MaxPG | unilateral | 125195\20247-3\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Mean PG | Mean PG | meas | m_p_MV_VTI_CardPGmean | unilateral | 125195\20256-4\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Minor | Length | meas | m_p_MV_Minor | unilateral | 125195\G-03A2\G-A194\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-A194, Minor Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV P1/2t | Pressure Half Time | meas | m_p_MV_PHalfT_PHalftime | unilateral | 125195\20280-4\T-35300 | SRT, T-35300, Mitral Valve | LN, 20280-4, Pressure Half-Time |  | UCUM, ms, Millisecond
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Peak A Vel | Unsigned Maximum Velocity | meas | m_p_MV_PeakAVel_Vel | unilateral | 125195\59081-0\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 59081-0, A-Wave Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Peak E Vel | Unsigned Maximum Velocity | meas | m_p_MV_PeakEVel_Vel | unilateral | 125195\59080-2\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 59080-2, E-Wave Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Radius | Length | meas | m_p_MV_Radius | unilateral | 125195\59102-4\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 59102-4, Flow Radius |  | UCUM, cm, Centimeter |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV R-R (DOP) | Time | meas | m_p_R_R_INT_MV | unilateral | 125195\8867-4\T-35300 | SRT, T-35300, Mitral Valve | LN, 8867-4, Heart Rate |  | UCUM, sec, Seconds
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Vmax | Unsigned Maximum Velocity | meas | m_p_MV_Vmax_Vmax | unilateral | 125195\20351-3\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV Vmean | Unsigned Time Average Peak Vel | meas | m_p_MV_VTI_CardTAPV | unilateral | 125195\11692-1\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 11692-1, Time Averaged Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV VTI | Velocity Time Integral | meas | m_p_MV_VTI_VTI | unilateral | 125195\20354-7\R-42047\T-35300 | SRT, T-35300, Mitral Valve | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MVA (P½t) | none | calc | c_p_MVA_PHalfT | unilateral | 125195\125210\G-038E\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | [5 empty] | DCM, 125210, Area by Pressure Half-Time
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MVA (PISA) | none | calc | c_p_MVA_PISA | unilateral | 125195\125216\G-038E\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | [5 empty] | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MVA (VTI) | none | calc | c_p_MVA_VTI | unilateral | 125195\125215\G-038E\T-35300 | SRT, T-35300, Mitral Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | [5 empty] | DCM, 125215, Continuity Equation by Velocity Time Integral
Pediatric Cardiac Ultrasound Report | Pediatric Echo | MV-HR (DOP) | Heart Rate | meas | m_p_HR_MV | unilateral | 125195\8867-4\T-35300 | SRT, T-35300, Mitral Valve | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, Beats Per Minute
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PA Acc Time | Time | meas | m_p_PA_AccelTime | unilateral | 125195\20168-1\T-44100 | SRT, T-44100, Pulmonary Trunk | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PAP (AT) | none | calc | c_p_PAP_AT | unilateral | 125195\59101-6\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 59101-6, Pulmonary Artery Pressure using Accel Time |  | UCUM, mm[Hg], Millimeters Of Mercury
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PDA Diam | Length | meas | m_p_PDA_Diam | unilateral | 125195\D4-32012\G-03A2\M-02550 | SRT, D4-32012, Patent Ductus Arteriosus | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PDA Dias Vel | Unsigned Maximum Velocity | meas | m_p_PDA_DiasVel_Vel | unilateral | 125195\20351-3\D4-32012\F-32011 | SRT, D4-32012, Patent Ductus Arteriosus | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PDA Sys Vel | Unsigned Maximum Velocity | meas | m_p_PDA_SysVel_Vel | unilateral | 125195\20351-3\D4-32012\R-FAB5B | SRT, D4-32012, Patent Ductus Arteriosus | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PDC Diam | Length | meas | m_p_PDC_Diam | unilateral | 125195\G-03A2\M-02550\T-D0878 | SRT, T-D0878, Posterior Descending Coronary Artery | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PI Decel Slope | Doppler Slope | meas | m_p_PI_DS_DopSlope | unilateral | 125195\20216-8\G-0367\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20216-8, Deceleration Slope |  | UCUM, cm/s2, Centimeter Per Second Square |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PI Decel Time | Time | meas | m_p_PI_DS_DopTime | unilateral | 125195\20217-6\G-0367\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20217-6, Deceleration Time |  | UCUM, sec, Seconds |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PI End Dias Vel | Unsigned Maximum Velocity | meas | m_p_PI_EndDiVel_Vel | unilateral | 125195\11653-3\G-0367\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PI P1/2t | Pressure Half Time | meas | m_p_PI_PHalfT_PHalftime | unilateral | 125195\20280-4\G-0367\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20280-4, Pressure Half-Time |  | UCUM, ms, Millisecond |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PI P1/2t Vmax | Unsigned Maximum Velocity | meas | m_p_PI_PHalfT_Vmax | unilateral | 125195\C12222-03\G-0367\T-35200 | SRT, T-35200, Pulmonic Valve | 99PMSBLUS, C12222-03, Pressure Half-Time Peak velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PISA (AI) | none | calc | c_p_PISA_AI | unilateral | 125195\125216\20226-7\G-0367\T-35400 | SRT, T-35400, Aortic Valve | LN, 20226-7, Flow Area |  | UCUM, cm2, Square Centimeter |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PISA (MR) | none | calc | c_p_PISA_MR | unilateral | 125195\125216\20226-7\G-0367\T-35300 | SRT, T-35300, Mitral Valve | LN, 20226-7, Flow Area |  | UCUM, cm2, Square Centimeter |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PISA (TR) | none | calc | c_p_PISA_TR | unilateral | 125195\125216\20226-7\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20226-7, Flow Area |  | UCUM, cm2, Square Centimeter |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Pulm V S/D | none | calc | c_p_PulmSD | unilateral | 125195\12144-2\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PulmV A Dur | Time | meas | m_p_PulmARevsDur | unilateral | 125195\59105-7\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 59105-7, A-Wave Duration |  | UCUM, sec, Seconds
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PulmV A Vel | Unsigned Maximum Velocity | meas | m_p_PulmARevsVel_Vel | unilateral | 125195\59079-4\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 59079-4, Peak Reversal Velocity during Atrial Contraction |  | UCUM, cm/s, Centimeter Per Second
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PulmV Dias Vel | Unsigned Maximum Velocity | meas | m_p_PulmDiasVel_Vel | unilateral | 125195\20351-3\F-32011\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PulmV Sys Vel | Unsigned Maximum Velocity | meas | m_p_PulmSysVel_Vel | unilateral | 125195\20351-3\R-FAB5B\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PV Acc Slope | Doppler Slope | meas | m_p_PV_AS_DopSlope | unilateral | 125195\20167-3\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20167-3, Acceleration Slope |  | UCUM, cm/s2, Centimeter Per Second Square |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PV Acc Time | Time | meas | m_p_PV_AS_DopTime | unilateral | 125195\20168-1\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PV Annul Diam | Length | meas | m_p_PV_Annul_Diam | unilateral | 125195\G-03A2\M-02550\T-3500E\T-35200 | SRT, T-35200, Pulmonic Valve | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter | SRT, T-3500E, Cardiac valve annulus | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PV Max PG | Maximum PG | meas | m_p_PV_Vmax_MaxPG | unilateral | 125195\20247-3\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PV Mean PG | Mean PG | meas | m_p_PV_VTI_CardPGmean | unilateral | 125195\20256-4\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PV R-R (DOP) | Time | meas | m_p_R_R_INT_PV | unilateral | 125195\8867-4\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 8867-4, Heart Rate |  | UCUM, sec, Seconds
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PV R-R (MM) | Time | meas | m_p_R_R_PV_MM | unilateral | 125195\8867-4\G-0394\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 8867-4, Heart Rate |  | UCUM, sec, Seconds |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PV Vmax | Unsigned Maximum Velocity | meas | m_p_PV_Vmax_Vmax | unilateral | 125195\20351-3\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PV Vmean | Unsigned Time Average Peak Vel | meas | m_p_PV_VTI_CardTAPV | unilateral | 125195\11692-1\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 11692-1, Time Averaged Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PV VTI | Velocity Time Integral | meas | m_p_PV_VTI_VTI | unilateral | 125195\20354-7\R-42047\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PVA (Vmax) | none | calc | c_p_PVA_Vmax | unilateral | 125195\125214\G-038E\T-35200 | SRT, T-35200, Pulmonic Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | [5 empty] | DCM, 125214, Continuity Equation by Peak Velocity
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PVA (VTI) | none | calc | c_p_PVA_VTI | unilateral | 125195\125215\G-038E\T-35200 | SRT, T-35200, Pulmonic Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | [5 empty] | DCM, 125215, Continuity Equation by Velocity Time Integral
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PV-HR (DOP) | Heart Rate | meas | m_p_HR_PV | unilateral | 125195\8867-4\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, Beats Per Minute
Pediatric Cardiac Ultrasound Report | Pediatric Echo | PV-HR (MM) | Heart Rate | meas | m_p_HR_PV_MM | unilateral | 125195\8867-4\G-0394\T-35200 | SRT, T-35200, Pulmonic Valve | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, Beats Per Minute |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Qp/Qs | none | calc | c_p_QpQs | unilateral | 125195\29462-9\P5-30031 | SRT, P5-30031, Cardiac Shunt Study | LN, 29462-9, Pulmonary-to-Systemic Shunt Flow Ratio |  | UCUM, 1, no units
Pediatric Cardiac Ultrasound Report | Pediatric Echo | R to AV Closure | Time | meas | m_p_RtoAVc_DopTime | unilateral | 125195\C12211-07\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12211-07, R Wave to Aortic Valve Closure Time |  | UCUM, ms, Millisecond
Pediatric Cardiac Ultrasound Report | Pediatric Echo | R to AV Open | Time | meas | m_p_RtoAVo_DopTime | unilateral | 125195\C12211-06\T-35400 | SRT, T-35400, Aortic Valve | 99PMSBLUS, C12211-06, R Wave to Aortic Valve Opening Time |  | UCUM, ms, Millisecond
Pediatric Cardiac Ultrasound Report | Pediatric Echo | R to MV Closure | Time | meas | m_p_RtoMVc_DopTime | unilateral | 125195\C12207-42\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-42, R Wave to Mitral Valve Closure Time |  | UCUM, ms, Millisecond
Pediatric Cardiac Ultrasound Report | Pediatric Echo | R to MV Open | Time | meas | m_p_RtoMVo_DopTime | unilateral | 125195\C12207-41\T-35300 | SRT, T-35300, Mitral Valve | 99PMSBLUS, C12207-41, R Wave to Mitral Valve Opening Time |  | UCUM, ms, Millisecond
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA Dimen (2D) | Length | meas | m_p_RA_Dimen_2D | unilateral | 125195\59090-1\G-03A2\T-32200 | SRT, T-32200, Right Atrium | LN, 59090-1, ROI Internal Dimension by US |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA Dimen (MM) | Length | meas | m_p_RA_Dimen_MM | unilateral | 125195\59090-1\G-0394\T-32200 | SRT, T-32200, Right Atrium | LN, 59090-1, ROI Internal Dimension by US |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RA Pressure | Pressure | meas | m_p_RA_Press_UserPressure | unilateral | 125195\18070-3\R-FAB5B\T-32200 | SRT, T-32200, Right Atrium | LN, 18070-3, Right Atrium Systolic Pressure |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RAED Major-A4C | Length | meas | m_p_RAED_MAJOR_A4C | unilateral | 125195\F-32011\G-03A2\G-A193\G-A19C\T-32200 | SRT, T-32200, Right Atrium | SRT, G-A193, Major Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RAED Minor-A4C | Length | meas | m_p_RAED_MINOR_A4C | unilateral | 125195\F-32011\G-03A2\G-A194\G-A19C\T-32200 | SRT, T-32200, Right Atrium | SRT, G-A194, Minor Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RAES Major-A4C | Length | meas | m_p_RAES_MAJOR_A4C | unilateral | 125195\G-03A2\G-A193\G-A19C\R-FAB5B\T-32200 | SRT, T-32200, Right Atrium | SRT, G-A193, Major Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RAES Minor-A4C | Length | meas | m_p_RAES_MINOR_A4C | unilateral | 125195\G-03A2\G-A194\G-A19C\R-FAB5B\T-32200 | SRT, T-32200, Right Atrium | SRT, G-A194, Minor Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RCA Diam | Length | meas | m_p_RCA_Diam | unilateral | 125195\G-03A2\M-02550\T-43203 | SRT, T-43203, Right Coronary Artery | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RL PulmV A Dur | Time | meas | m_p_RL_PulmARevsDur | unilateral | 125195\59105-7\T-48520\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 59105-7, A-Wave Duration |  | UCUM, sec, Seconds | SRT, T-48520, Right Inferior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RL PulmV A Vel | Unsigned Maximum Velocity | meas | m_p_RL_PulmARevsVel_Vel | unilateral | 125195\59079-4\T-48520\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 59079-4, Peak Reversal Velocity during Atrial Contraction |  | UCUM, cm/s, Centimeter Per Second | SRT, T-48520, Right Inferior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RL PulmV Diam | Length | meas | m_p_RL_PulmV_Diam | unilateral | 125195\G-03A2\M-02550\T-48520\T-48581 | SRT, T-48581, Pulmonary Vein | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter | SRT, T-48520, Right Inferior Pulmonary Vein | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RL PulmV Dias Vel | Unsigned Maximum Velocity | meas | m_p_RL_PulmDiasVel_Vel | unilateral | 125195\20351-3\F-32011\T-48520\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, T-48520, Right Inferior Pulmonary Vein |  |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RL PulmV S/D | none | calc | c_p_RL_PulmSD | unilateral | 125195\12144-2\T-48520\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | SRT, T-48520, Right Inferior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RL PulmV Sys Vel | Unsigned Maximum Velocity | meas | m_p_RL_PulmSysVel_Vel | unilateral | 125195\20351-3\R-FAB5B\T-48520\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, T-48520, Right Inferior Pulmonary Vein |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RPA Diam | Length | meas | m_p_RPA_Diam | unilateral | 125195\G-03A2\M-02550\T-44000\T-44200 | SRT, T-44000, Pulmonary artery | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter | SRT, T-44200, Right pulmonary artery | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RPA Max PG | Maximum PG | meas | m_p_RPA_Vmax_MaxPG | unilateral | 125195\20247-3\T-44000\T-44200 | SRT, T-44000, Pulmonary artery | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury | SRT, T-44200, Right pulmonary artery
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RPA Mean PG | Mean PG | meas | m_p_RPA_PG | unilateral | 125195\20256-4\T-44000\T-44200 | SRT, T-44000, Pulmonary artery | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury | SRT, T-44200, Right pulmonary artery
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RPA Vmax | Unsigned Maximum Velocity | meas | m_p_RPA_Vmax_Vmax | unilateral | 125195\20351-3\T-44000\T-44200 | SRT, T-44000, Pulmonary artery | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, T-44200, Right pulmonary artery
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RU PulmV A Dur | Time | meas | m_p_RU_PulmARevsDur | unilateral | 125195\59105-7\T-48510\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 59105-7, A-Wave Duration |  | UCUM, sec, Seconds | SRT, T-48510, Right Superior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RU PulmV A Vel | Unsigned Maximum Velocity | meas | m_p_RU_PulmARevsVel_Vel | unilateral | 125195\59079-4\T-48510\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 59079-4, Peak Reversal Velocity during Atrial Contraction |  | UCUM, cm/s, Centimeter Per Second | SRT, T-48510, Right Superior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RU PulmV Diam | Length | meas | m_p_RU_PulmV_Diam | unilateral | 125195\G-03A2\M-02550\T-48510\T-48581 | SRT, T-48581, Pulmonary Vein | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter | SRT, T-48510, Right Superior Pulmonary Vein | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RU PulmV Dias Vel | Unsigned Maximum Velocity | meas | m_p_RU_PulmDiasVel_Vel | unilateral | 125195\20351-3\F-32011\T-48510\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, T-48510, Right Superior Pulmonary Vein |  |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RU PulmV S/D | none | calc | c_p_RU_PulmSD | unilateral | 125195\12144-2\T-48510\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units | SRT, T-48510, Right Superior Pulmonary Vein
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RU PulmV Sys Vel | Unsigned Maximum Velocity | meas | m_p_RU_PulmSysVel_Vel | unilateral | 125195\20351-3\R-FAB5B\T-48510\T-48581 | SRT, T-48581, Pulmonary Vein | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, T-48510, Right Superior Pulmonary Vein |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RV ET (DOP) | Time | meas | m_p_RV_ET_Dop | unilateral | 125195\20222-6\T-32500 | SRT, T-32500, Right Ventricle | LN, 20222-6, Ejection Time |  | UCUM, sec, Seconds
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RV ET (MM) | Time | meas | m_p_RV_Eject_Tm | unilateral | 125195\20222-6\G-0394\T-32500 | SRT, T-32500, Right Ventricle | LN, 20222-6, Ejection Time |  | UCUM, sec, Seconds |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RV MPI | none | calc | c_p_RV_MPI | unilateral | 125195\59099-2\T-32500 | SRT, T-32500, Right Ventricle | LN, 59099-2, Myocardial Performance Index (Tei) |  | UCUM, 1, no units
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RV PEP | Time | meas | m_p_RV_PreEject_Tm | unilateral | 125195\59085-1\G-0394\T-32500 | SRT, T-32500, Right Ventricle | LN, 59085-1, Pre-Ejection Period |  | UCUM, sec, Seconds |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RV PEP/ET | none | calc | c_p_RV_PEP_ET | unilateral | 125195\59088-5\G-0394\T-32500 | SRT, T-32500, Right Ventricle | LN, 59088-5, Pre-Ejection Period/Ejection Time Ratio |  | UCUM, 1, no units |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RV PEPc | none | calc | c_p_RVPEPc | unilateral | 125195\59087-7\G-0394\T-32500 | SRT, T-32500, Right Ventricle | LN, 59087-7, Heart Rate-Corrected Pre-Ejection Period |  | UCUM, ms, Millisecond |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVAWd (2D) | Length | meas | m_p_RVAWd_2D | unilateral | 125195\59089-3\F-32011\G-03A2\R-4210B\T-32500 | SRT, T-32500, Right Ventricle | LN, 59089-3, ROI Thickness by US |  | UCUM, cm, Centimeter | SRT, R-4210B, Anterior Wall | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVAWd (MM) | Length | meas | m_p_RVAWd_MM | unilateral | 125195\59089-3\F-32011\G-0394\R-4210B\T-32500 | SRT, T-32500, Right Ventricle | LN, 59089-3, ROI Thickness by US |  | UCUM, cm, Centimeter | SRT, R-4210B, Anterior Wall | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVED Major-A4C | Length | meas | m_p_RVED_MAJOR_A4C | unilateral | 125195\F-32011\G-03A2\G-A193\G-A19C\T-32500 | SRT, T-32500, Right Ventricle | SRT, G-A193, Major Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVED Minor-A4C | Length | meas | m_p_RVED_MINOR_A4C | unilateral | 125195\F-32011\G-03A2\G-A194\G-A19C\T-32500 | SRT, T-32500, Right Ventricle | SRT, G-A194, Minor Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVES Major-A4C | Length | meas | m_p_RVES_MAJOR_A4C | unilateral | 125195\G-03A2\G-A193\G-A19C\R-FAB5B\T-32500 | SRT, T-32500, Right Ventricle | SRT, G-A193, Major Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVES Minor-A4C | Length | meas | m_p_RVES_MINOR_A4C | unilateral | 125195\G-03A2\G-A194\G-A19C\R-FAB5B\T-32500 | SRT, T-32500, Right Ventricle | SRT, G-A194, Minor Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVIDd (2D) | Length | meas | m_p_RVIDd_2D | unilateral | 125195\59090-1\F-32011\G-03A2\T-32500 | SRT, T-32500, Right Ventricle | LN, 59090-1, ROI Internal Dimension by US |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVIDd (MM) | Length | meas | m_p_RVIDd_MM | unilateral | 125195\59090-1\F-32011\G-0394\T-32500 | SRT, T-32500, Right Ventricle | LN, 59090-1, ROI Internal Dimension by US |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVOT Acc Slope | Doppler Slope | meas | m_p_RVOT_AS_DopSlope | unilateral | 125195\20167-3\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20167-3, Acceleration Slope |  | UCUM, cm/s2, Centimeter Per Second Square | SRT, T-32550, Right Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVOT Acc Time | Time | meas | m_p_RVOT_AS_DopTime | unilateral | 125195\20168-1\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | SRT, T-32550, Right Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVOT Area | none | calc | c_p_RVOT_Area | unilateral | 125195\G-038E\G-03A2\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | SRT, T-32550, Right Ventricle Outflow Tract | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVOT Diam | Length | meas | m_p_RVOT_Diam | unilateral | 125195\G-038F\G-03A2\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, G-038F, Cardiovascular Orifice Diameter |  | UCUM, cm, Centimeter | SRT, T-32550, Right Ventricle Outflow Tract | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVOT Max PG | Maximum PG | meas | m_p_RVOT_Vmax_MaxPG | unilateral | 125195\20247-3\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury | SRT, T-32550, Right Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVOT Mean PG | Mean PG | meas | m_p_RVOT_VTI_CardPGmean | unilateral | 125195\20256-4\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury | SRT, T-32550, Right Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVOT Vmax | Unsigned Maximum Velocity | meas | m_p_RVOT_Vmax_Vmax | unilateral | 125195\20351-3\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, T-32550, Right Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVOT Vmean | Unsigned Time Average Peak Vel | meas | m_p_RVOT_VTI_CardTAPV | unilateral | 125195\11692-1\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 11692-1, Time Averaged Peak Velocity |  | UCUM, cm/s, Centimeter Per Second | SRT, T-32550, Right Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVOT VTI | Velocity Time Integral | meas | m_p_RVOT_VTI_VTI | unilateral | 125195\20354-7\T-32500\T-32550 | SRT, T-32500, Right Ventricle | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | SRT, T-32550, Right Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVSP | none | calc | c_p_RVSP | unilateral | 125195\G-0380\R-FAB5B\T-32500 | SRT, T-32500, Right Ventricle | SRT, G-0380, Right Ventricular Peak Systolic Pressure |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | RVSP (VSD) | none | calc | c_p_RVSP_VSD | unilateral | 125195\D4-31150\F-31000\R-FAB5B | SRT, D4-31150, Ventricular Septal Defect | SRT, F-31000, Blood Pressure |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SI (2D-Cubed) | none | calc | c_p_SI_Cubed | unilateral | 125195\125206\F-00078\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SI (2D-Teich) | none | calc | c_p_SI_Teich | unilateral | 125195\125209\F-00078\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SI (A/L) | none | calc | c_p_SI_SPE | unilateral | 125195\125226\F-00078\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SI (A2C) | none | calc | c_p_SI_MOD_A2C | unilateral | 125195\125208\F-00078\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SI (A2C-A/L) | none | calc | c_p_SI_A2C_AL | unilateral | 125195\125226\F-00078\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SI (A4C) | none | calc | c_p_SI_MOD_A4C | unilateral | 125195\125208\F-00078\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SI (A4C-A/L) | none | calc | c_p_SI_A4C_AL | unilateral | 125195\125226\F-00078\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SI (BP) | none | calc | c_p_SI_MOD_BP | unilateral | 125195\125207\F-00078\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  | UCUM, ml/m2, ml/m2 |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SI (MM-Cubed) | none | calc | c_p_SI_Cubed_MM | unilateral | 125195\125206\F-00078\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  | UCUM, ml/m2, ml/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SI (MM-Teich) | none | calc | c_p_SI_Teich_MM | unilateral | 125195\125209\F-00078\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-00078, Stroke Index |  | UCUM, ml/m2, ml/m2 |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (2D-Cubed) | none | calc | c_p_SV_Cubed | unilateral | 125195\125206\F-32120\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (2D-Teich) | none | calc | c_p_SV_Teich | unilateral | 125195\125209\F-32120\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (A/L) | none | calc | c_p_SV_SPE | unilateral | 125195\125226\F-32120\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (A2C) | none | calc | c_p_SV_MOD_A2C | unilateral | 125195\125208\F-32120\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (A2C-A/L) | none | calc | c_p_SV_A2C_AL | unilateral | 125195\125226\F-32120\G-03A2\G-A19B\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19B, Apical two chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (A4C) | none | calc | c_p_SV_MOD_A4C | unilateral | 125195\125208\F-32120\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125208, Method of Disks, Single Plane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (A4C-A/L) | none | calc | c_p_SV_A4C_AL | unilateral | 125195\125226\F-32120\G-03A2\G-A19C\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode | SRT, G-A19C, Apical four chamber |  |  | DCM, 125226, Single Plane Ellipse
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (BP) | none | calc | c_p_SV_MOD_BP | unilateral | 125195\125207\F-32120\G-03A2\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter |  | SRT, G-03A2, 2D mode |  |  |  | DCM, 125207, Method of Disks, Biplane
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (LVOT) | none | calc | c_p_SV_LVOT | unilateral | 125195\F-32120\T-32600\T-32650 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter | SRT, T-32650, Left Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (MM-Cubed) | none | calc | c_p_SV_Cubed_MM | unilateral | 125195\125206\F-32120\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter |  | SRT, G-0394, M mode |  |  |  | DCM, 125206, Cube Method
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (MM-Teich) | none | calc | c_p_SV_Teich_MM | unilateral | 125195\125209\F-32120\G-0394\T-32600 | SRT, T-32600, Left Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter |  | SRT, G-0394, M mode |  |  |  | DCM, 125209, Teichholz
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (MV) | none | calc | c_p_SV_MV | unilateral | 125195\F-32120\T-35300 | SRT, T-35300, Mitral Valve | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (PV) | none | calc | c_p_SV_PV | unilateral | 125195\F-32120\T-35200 | SRT, T-35200, Pulmonic Valve | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (RVOT) | none | calc | c_p_SV_RVOT | unilateral | 125195\F-32120\T-32500\T-32550 | SRT, T-32500, Right Ventricle | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter | SRT, T-32550, Right Ventricle Outflow Tract
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SV (TV) | none | calc | c_p_SV_TV | unilateral | 125195\F-32120\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, F-32120, Stroke Volume |  | UCUM, ml, Milliliter
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SVC A Dur | Time | meas | m_p_SVC_ARevsDur | unilateral | 125195\59105-7\M-2460D | SRT, M-2460D, Right Superior vena cava | LN, 59105-7, A-Wave Duration |  | UCUM, sec, Seconds
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SVC A Vel | Unsigned Maximum Velocity | meas | m_p_SVC_ARevsVel_Vel | unilateral | 125195\59079-4\M-2460D | SRT, M-2460D, Right Superior vena cava | LN, 59079-4, Peak Reversal Velocity during Atrial Contraction |  | UCUM, cm/s, Centimeter Per Second
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SVC Diam | Length | meas | m_p_SVC_Diam | unilateral | 125195\G-03A2\M-02550\M-2460D | SRT, M-2460D, Right Superior vena cava | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SVC Dias Vel | Unsigned Maximum Velocity | meas | m_p_SVC_DiasVel_Vel | unilateral | 125195\20351-3\F-32011\M-2460D | SRT, M-2460D, Right Superior vena cava | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  | SRT, F-32011, End Diastole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SVC S/D | none | calc | c_p_SVC_SD | unilateral | 125195\12144-2\M-2460D | SRT, M-2460D, Right Superior vena cava | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, no units
Pediatric Cardiac Ultrasound Report | Pediatric Echo | SVC Sys Vel | Unsigned Maximum Velocity | meas | m_p_SVC_SysVel_Vel | unilateral | 125195\20351-3\M-2460D\R-FAB5B | SRT, M-2460D, Right Superior vena cava | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  | SRT, R-FAB5B, End Systole
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Tei Index | none | calc | c_p_TeiIndex | unilateral | 125195\59099-2\T-35300 | SRT, T-35300, Mitral Valve | LN, 59099-2, Myocardial Performance Index (Tei) |  | UCUM, 1, no units
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Time to Lat E` | Time | meas | m_p_TimeLatE | unilateral | 125195\59096-8\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59096-8, Time to Left Ventricle E Tissue Velocity |  | UCUM, sec, Seconds | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Time to Lat S | Time | meas | m_p_TimeLatS | unilateral | 125195\59095-0\G-0392\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59095-0, Time to Left Ventricle S Tissue Velocity |  | UCUM, sec, Seconds | SRT, G-0392, Lateral Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Time to Med E` | Time | meas | m_p_TimeMedE | unilateral | 125195\59096-8\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59096-8, Time to Left Ventricle E Tissue Velocity |  | UCUM, sec, Seconds | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Time to Med S | Time | meas | m_p_TimeMedS | unilateral | 125195\59095-0\G-0391\P5-B0128\T-32600 | SRT, T-32600, Left Ventricle | LN, 59095-0, Time to Left Ventricle S Tissue Velocity |  | UCUM, sec, Seconds | SRT, G-0391, Medial Mitral Annulus | SRT, P5-B0128, Tissue Doppler Imaging
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TR Alias Vel | Unsigned Maximum Velocity | meas | m_p_TR_AliasVel | unilateral | 125195\59130-5\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59130-5, Alias velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TR ERO | none | calc | c_p_TR_ERO | unilateral | 125195\125216\G-0367\G-038E\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TR Flow Rate | none | calc | c_p_TR_FlowRate | unilateral | 125195\34141-2\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 34141-2, Peak Instantaneous Flow Rate |  | UCUM, ml/sec, ml/sec |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TR Fraction | none | calc | c_p_TR_Frac | unilateral | 125195\G-0367\G-0390\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-0390, Regurgitant Fraction |  | UCUM, %, Percent |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TR Max PG | Maximum PG | meas | m_p_TR_Vmax_MaxPG | unilateral | 125195\20247-3\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TR Mean PG | Mean PG | meas | m_p_TR_VTI_CardPGmean | unilateral | 125195\20256-4\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TR Radius | Length | meas | m_p_TR_Radius | unilateral | 125195\59102-4\G-0367\G-03A2\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59102-4, Flow Radius |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TR Vmax | Unsigned Maximum Velocity | meas | m_p_TR_Vmax_Vmax | unilateral | 125195\20351-3\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TR Vmean | Unsigned Time Average Peak Vel | meas | m_p_TR_VTI_CardTAPV | unilateral | 125195\11692-1\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 11692-1, Time Averaged Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TR Volume | none | calc | c_p_TR_Vol | unilateral | 125195\125216\33878-0\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 33878-0, Volume Flow |  | UCUM, ml, Milliliter |  |  |  |  | SRT, G-0367, Regurgitant Flow | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TR VTI | Velocity Time Integral | meas | m_p_TR_VTI_VTI | unilateral | 125195\20354-7\G-0367\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter |  |  |  |  | SRT, G-0367, Regurgitant Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV A Dur | Time | meas | m_p_TV_A_Dur | unilateral | 125195\59105-7\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59105-7, A-Wave Duration |  | UCUM, sec, Seconds
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV A-C Interval | Time | meas | m_p_TV_AC_Int | unilateral | 125195\59103-2\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59103-2, A-C Interval |  | UCUM, sec, Seconds |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Acc Slope | Doppler Slope | meas | m_p_TV_AS_DopSlope | unilateral | 125195\20167-3\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20167-3, Acceleration Slope |  | UCUM, cm/s2, Centimeter Per Second Square |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Acc Time | Time | meas | m_p_TV_AS_DopTime | unilateral | 125195\20168-1\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Alias Vel | Unsigned Maximum Velocity | meas | m_p_TV_AliasVel | unilateral | 125195\59130-5\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59130-5, Alias velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Annul Area | none | calc | c_p_TV_Annul_Area | unilateral | 125195\G-038E\G-03A2\T-3500E\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | SRT, T-3500E, Cardiac valve annulus | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Annul Diam | Length | meas | m_p_TV_Annul_Diam | unilateral | 125195\G-03A2\M-02550\T-3500E\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter | SRT, T-3500E, Cardiac valve annulus | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Area | none | calc | c_p_TV_Area | unilateral | 125195\G-03A2\G-A166\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-A166, Area |  | UCUM, cm2, Square Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Closure to Opening | Time | meas | m_p_TV_Cl_to_Op | unilateral | 125195\59082-8\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59082-8, Closure to Opening Time |  | UCUM, sec, Seconds
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV D-E Exc Dist | Length | meas | m_p_TV_DE_Exc_Dist | unilateral | 125195\59091-9\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59091-9, D-E Excursion |  | UCUM, cm, Centimeter |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV D-E Slope | MMode Slope | meas | m_p_TV_DE_Slope_MMSlope | unilateral | 125195\59127-1\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59127-1, D-E Slope |  | UCUM, cm/s, Centimeter Per Second |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Decel Slope | Doppler Slope | meas | m_p_TV_DS_DopSlope | unilateral | 125195\20216-8\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20216-8, Deceleration Slope |  | UCUM, cm/s2, Centimeter Per Second Square |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Decel Time | Time | meas | m_p_TV_DS_DopTime | unilateral | 125195\20217-6\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20217-6, Deceleration Time |  | UCUM, sec, Seconds |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV DFP | Time | meas | m_p_TV_DFP | unilateral | 125195\R-003A9\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, R-003A9, Tricuspid Diastolic Filling Period (DFPt) |  | UCUM, sec, Seconds
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Diam | Length | meas | m_p_TV_Diam | unilateral | 125195\G-038F\G-03A2\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-038F, Cardiovascular Orifice Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV E/A | none | calc | c_p_TV_EA | unilateral | 125195\59104-0\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59104-0, Peak E wave/Peak A wave by US |  | UCUM, 1, no units
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV E-F Slope | MMode Slope | meas | m_p_TV_EF_Slope_MMSlope | unilateral | 125195\59128-9\G-0394\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59128-9, E-F Slope |  | UCUM, cm/s, Centimeter Per Second |  | SRT, G-0394, M mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Max PG | Maximum PG | meas | m_p_TV_Vmax_MaxPG | unilateral | 125195\20247-3\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Mean PG | Mean PG | meas | m_p_TV_VTI_CardPGmean | unilateral | 125195\20256-4\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20256-4, Mean Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV P1/2t | Pressure Half Time | meas | m_p_TV_PHalfT_PHalftime | unilateral | 125195\20280-4\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20280-4, Pressure Half-Time |  | UCUM, ms, Millisecond
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV P1/2t Vmax | Unsigned Maximum Velocity | meas | m_p_TV_PHalfT_Vmax | unilateral | 125195\C12222-03\T-35100 | SRT, T-35100, Tricuspid Valve | 99PMSBLUS, C12222-03, Pressure Half-Time Peak velocity |  | UCUM, cm/s, Centimeter Per Second
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Peak A Vel | Unsigned Maximum Velocity | meas | m_p_TV_PeakAVel_Vel | unilateral | 125195\59081-0\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59081-0, A-Wave Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Peak E Vel | Unsigned Maximum Velocity | meas | m_p_TV_PeakEVel_Vel | unilateral | 125195\59080-2\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59080-2, E-Wave Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Radius | Length | meas | m_p_TV_Radius | unilateral | 125195\59102-4\G-03A2\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 59102-4, Flow Radius |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV R-R (DOP) | Time | meas | m_p_R_R_INT_TV | unilateral | 125195\8867-4\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 8867-4, Heart Rate |  | UCUM, sec, Seconds
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Vmax | Unsigned Maximum Velocity | meas | m_p_TV_Vmax_Vmax | unilateral | 125195\20351-3\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV Vmean | Unsigned Time Average Peak Vel | meas | m_p_TV_VTI_CardTAPV | unilateral | 125195\11692-1\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 11692-1, Time Averaged Peak Velocity |  | UCUM, cm/s, Centimeter Per Second |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV VTI | Velocity Time Integral | meas | m_p_TV_VTI_VTI | unilateral | 125195\20354-7\R-42047\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter |  |  |  |  | SRT, R-42047, Antegrade Flow
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TVA (PISA) | none | calc | c_p_TVA_PISA | unilateral | 125195\125216\G-038E\T-35100 | SRT, T-35100, Tricuspid Valve | SRT, G-038E, Cardiovascular Orifice Area |  | UCUM, cm2, Square Centimeter | [5 empty] | DCM, 125216, Proximal Isovelocity Surface Area
Pediatric Cardiac Ultrasound Report | Pediatric Echo | TV-HR (DOP) | Heart Rate | meas | m_p_HR_TV | unilateral | 125195\8867-4\T-35100 | SRT, T-35100, Tricuspid Valve | LN, 8867-4, Heart Rate |  | UCUM, {H.B.}/min, Beats Per Minute
Pediatric Cardiac Ultrasound Report | Pediatric Echo | VSD Diam | none | calc | c_p_VSD_Diam | unilateral | 125195\D4-31150\G-03A2\M-02550 | SRT, D4-31150, Ventricular Septal Defect | SRT, M-02550, Diameter |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | VSD Major | Length | meas | m_p_VSD_Major | unilateral | 125195\D4-31150\G-03A2\G-A193 | SRT, D4-31150, Ventricular Septal Defect | SRT, G-A193, Major Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | VSD Max PG | Maximum PG | meas | m_p_VSD_Vmax_MaxPG | unilateral | 125195\20247-3\D4-31150 | SRT, D4-31150, Ventricular Septal Defect | LN, 20247-3, Peak Gradient |  | UCUM, mm[Hg], Millimeters Of Mercury
Pediatric Cardiac Ultrasound Report | Pediatric Echo | VSD Minor | Length | meas | m_p_VSD_Minor | unilateral | 125195\D4-31150\G-03A2\G-A194 | SRT, D4-31150, Ventricular Septal Defect | SRT, G-A194, Minor Axis |  | UCUM, cm, Centimeter |  | SRT, G-03A2, 2D mode
Pediatric Cardiac Ultrasound Report | Pediatric Echo | VSD Vmax | Unsigned Maximum Velocity | meas | m_p_VSD_Vmax_Vmax | unilateral | 125195\20351-3\D4-31150 | SRT, D4-31150, Ventricular Septal Defect | LN, 20351-3, Peak Velocity |  | UCUM, cm/s, Centimeter Per Second
Pediatric Cardiac Ultrasound Report | Pediatric Echo | Wall Stress | none | calc | c_p_Wall_Stress | unilateral | 125195\59097-6\G-0394\T-32600 | SRT, T-32600, Left Ventricle | LN, 59097-6, Left Ventricle Meridional Wall Stress |  | UCUM, kg/m2, kg/m2 |  | SRT, G-0394, M mode
Vascular Ultrasound Procedure Report | Vascular | ACA EDV | Signed End Diastolic Vel | meas | m_tcd_ACA_A1_EDV | bilateral | 125100\11653-3\T-40501\T-45540 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45540, Anterior Cerebral Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | ACA PI | none | calc | c_tcd_ACA_A1_PI | bilateral | 125100\12008-9\T-40501\T-45540 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45540, Anterior Cerebral Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | ACA PSV | Signed Maximum Velocity | meas | m_tcd_ACA_A1_PSV | bilateral | 125100\11726-7\T-40501\T-45540 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45540, Anterior Cerebral Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | ACA TAPV | Signed Time Average Peak Vel | meas | m_tcd_ACA_A1_TAPV | bilateral | 125100\11692-1\T-40501\T-45540 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45540, Anterior Cerebral Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | ACA TCD Mean | none | calc | c_tcd_ACA_A1_TcdMean | bilateral | 125100\C12221-03\T-40501\T-45540 |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45540, Anterior Cerebral Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | ACA TCD PI | none | calc | c_tcd_ACA_A1_TcdPI | bilateral | 125100\C12221-04\T-40501\T-45540 |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45540, Anterior Cerebral Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | ACoA EDV | Signed End Diastolic Vel | meas | m_tcd_AComA_EDV | bilateral | 125100\11653-3\T-40501\T-45530 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45530, Anterior Communicating Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | ACoA PI | none | calc | c_tcd_AComA_PI | bilateral | 125100\12008-9\T-40501\T-45530 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45530, Anterior Communicating Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | ACoA PSV | Signed Maximum Velocity | meas | m_tcd_AComA_PSV | bilateral | 125100\11726-7\T-40501\T-45530 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45530, Anterior Communicating Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | ACoA TAPV | Signed Time Average Peak Vel | meas | m_tcd_AComA_TAPV | bilateral | 125100\11692-1\T-40501\T-45530 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45530, Anterior Communicating Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | ACoA TCD Mean | none | calc | c_tcd_AComA_TcdMean | bilateral | 125100\C12221-03\T-40501\T-45530 |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45530, Anterior Communicating Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | ACoA TCD PI | none | calc | c_tcd_AComA_TcdPI | bilateral | 125100\C12221-04\T-40501\T-45530 |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45530, Anterior Communicating Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Antecube_AS | bilateral | 125100\20167-3\sup71_001\T-47020 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | 99PMSBLUS, sup71_001, Antecube |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Axillary_A_AS | bilateral | 125100\20167-3\T-47020\T-47100 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47100, Axillary Artery |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Brach_Ceph_A_AS | bilateral | 125100\20167-3\T-46010\T-47020 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46010, Innominate Artery |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Bulb_AS | bilateral | 125100\20167-3\T-45005\T-45170 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-45170, Carotid Bulb |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_CFA_AS | bilateral | 125100\20167-3\T-47040\T-47400 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47400, Common Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Com_Iliac_A_AS | bilateral | 125100\20167-3\T-46710\T-47040 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46710, Common Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Dist_ATA_AS | bilateral | 125100\20167-3\G-A119\T-47040\T-47700 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Dist_Brachial_A_AS | bilateral | 125100\20167-3\G-A119\T-47020\T-47160 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47160, Brachial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Dist_CCA_AS | bilateral | 125100\20167-3\G-A119\T-45005\T-45100 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Dist_ICA_AS | bilateral | 125100\20167-3\G-A119\T-45005\T-45300 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Dist_Pero_A_AS | bilateral | 125100\20167-3\G-A119\T-47040\T-47630 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Dist_POPA_AS | bilateral | 125100\20167-3\G-A119\T-47040\T-47500 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47500, Popliteal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Dist_PTA_AS | bilateral | 125100\20167-3\G-A119\T-47040\T-47600 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Dist_Radial_A_AS | bilateral | 125100\20167-3\G-A119\T-47020\T-47300 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Dist_SFA_AS | bilateral | 125100\20167-3\G-A119\T-47040\T-47403 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Dist_Ulnar_A_AS | bilateral | 125100\20167-3\G-A119\T-47020\T-47200 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Dor_Ped_ATA_AS | bilateral | 125100\20167-3\T-47040\T-47741 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47741, Dorsalis Pedis Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Ext_Iliac_A_AS | bilateral | 125100\20167-3\T-46910\T-47040 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46910, External Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Int_Iliac_A_AS | bilateral | 125100\20167-3\T-46740\T-47040 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46740, Internal Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Mid_ATA_AS | bilateral | 125100\20167-3\G-A188\T-47040\T-47700 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Mid_CCA_AS | bilateral | 125100\20167-3\G-A188\T-45005\T-45100 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Mid_ECA_AS | bilateral | 125100\20167-3\G-A188\T-45005\T-45200 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-45200, External Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Mid_ICA_AS | bilateral | 125100\20167-3\G-A188\T-45005\T-45300 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Mid_Pero_A_AS | bilateral | 125100\20167-3\G-A188\T-47040\T-47630 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Mid_PTA_AS | bilateral | 125100\20167-3\G-A188\T-47040\T-47600 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Mid_Radial_A_AS | bilateral | 125100\20167-3\G-A188\T-47020\T-47300 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Mid_SCL_A_AS | bilateral | 125100\20167-3\G-A188\T-45005\T-46100 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46100, Subclavian Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Mid_SFA_AS | bilateral | 125100\20167-3\G-A188\T-47040\T-47403 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Mid_Ulnar_A_AS | bilateral | 125100\20167-3\G-A188\T-47020\T-47200 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_PFA_AS | bilateral | 125100\20167-3\G-A118\T-47040\T-47440 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47440, Profunda Femoris Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Prox_ATA_AS | bilateral | 125100\20167-3\G-A118\T-47040\T-47700 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Prox_Brachial_A_AS | bilateral | 125100\20167-3\G-A118\T-47020\T-47160 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47160, Brachial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Prox_CCA_AS | bilateral | 125100\20167-3\G-A118\T-45005\T-45100 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Prox_ECA_AS | bilateral | 125100\20167-3\G-A118\T-45005\T-45200 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-45200, External Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Prox_ICA_AS | bilateral | 125100\20167-3\G-A118\T-45005\T-45300 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Prox_Pero_A_AS | bilateral | 125100\20167-3\G-A118\T-47040\T-47630 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Prox_POPA_AS | bilateral | 125100\20167-3\G-A118\T-47040\T-47500 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47500, Popliteal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Prox_PTA_AS | bilateral | 125100\20167-3\G-A118\T-47040\T-47600 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Prox_Radial_A_AS | bilateral | 125100\20167-3\G-A118\T-47020\T-47300 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Prox_SCL_A_AS | bilateral | 125100\20167-3\G-A118\T-45005\T-46100 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46100, Subclavian Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Prox_SFA_AS | bilateral | 125100\20167-3\G-A118\T-47040\T-47403 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Prox_Ulnar_A_AS | bilateral | 125100\20167-3\G-A118\T-47020\T-47200 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AI | Doppler Slope | meas | m_Vertebral_A_AS | bilateral | 125100\20167-3\T-45005\T-45700 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-45700, Vertebral Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | Antecube PI | none | calc | c_Antecube_PI | bilateral | 125100\12008-9\sup71_001\T-47020 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | 99PMSBLUS, sup71_001, Antecube |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | Antecube RI | none | calc | c_Antecube_RI | bilateral | 125100\12023-8\sup71_001\T-47020 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | 99PMSBLUS, sup71_001, Antecube |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | Antecube S/D | none | calc | c_Antecube_SD | bilateral | 125100\12144-2\sup71_001\T-47020 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | 99PMSBLUS, sup71_001, Antecube |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | Ao % Area Reduction | none | calc | c_Ao_Area_Reduc | unilateral | 125100\G-0371\T-42000\T-46002 |  | SRT, G-0371, Percent Area Reduction |  | UCUM, %, Percent | [6 empty] | SRT, T-42000, Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Ao % Diam Reduction | none | calc | c_Ao_Diam_Reduc | unilateral | 125100\G-0372\T-42000\T-46002 |  | SRT, G-0372, Percent Diameter Reduction |  | UCUM, %, Percent | [6 empty] | SRT, T-42000, Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Ao Area Resid | Area | meas | m_Ao_Area_Resid | unilateral | 125100\C7471-01\T-42000\T-46002 |  | 99PMSBLUS, C7471-01, Area 1 of Area Percent Reduction |  | UCUM, cm2, Square Centimeter | [6 empty] | SRT, T-42000, Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Ao Area True | Area | meas | m_Ao_Area_True | unilateral | 125100\C7471-02\T-42000\T-46002 |  | 99PMSBLUS, C7471-02, Area 2 of Area Percent Reduction |  | UCUM, cm2, Square Centimeter | [6 empty] | SRT, T-42000, Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Ao Diam Resid | Length | meas | m_Ao_Diam_Resid | unilateral | 125100\C7470-05\T-42000\T-46002 |  | 99PMSBLUS, C7470-05, Diameter 1 of Diameter Reduction |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-42000, Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Ao Diam True | Length | meas | m_Ao_Diam_True | unilateral | 125100\C7470-06\T-42000\T-46002 |  | 99PMSBLUS, C7470-06, Diameter 2 of Diameter Reduction |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-42000, Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Ao Dist Diam | Length | meas | m_aoDist | unilateral | 125100\G-0364\G-A119\T-42000\T-46002 |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-42000, Aorta |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Ao Mid Diam | Length | meas | m_aoMid | unilateral | 125100\G-0364\G-A188\T-42000\T-46002 |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-42000, Aorta |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Ao Prox Diam | Length | meas | m_aoProx | unilateral | 125100\G-0364\G-A118\T-42000\T-46002 |  | SRT, G-0364, Vessel lumen diameter |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-42000, Aorta |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Aorta PSV | Signed Maximum Velocity | meas | m_abdren_Ao_PSV | unilateral | 125100\11726-7\T-42000\T-46002 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-42000, Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Aorta PSV | Signed Maximum Velocity | meas | m_Ao_PSV | unilateral | 125100\11726-7\T-42000\T-46002 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-42000, Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Antecube_AT | bilateral | 125100\20168-1\sup71_001\T-47020 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | 99PMSBLUS, sup71_001, Antecube |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Axillary_A_AT | bilateral | 125100\20168-1\T-47020\T-47100 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47100, Axillary Artery |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Brach_Ceph_A_AT | bilateral | 125100\20168-1\T-46010\T-47020 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46010, Innominate Artery |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Bulb_AT | bilateral | 125100\20168-1\T-45005\T-45170 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-45170, Carotid Bulb |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_CFA_AT | bilateral | 125100\20168-1\T-47040\T-47400 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47400, Common Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Com_Iliac_A_AT | bilateral | 125100\20168-1\T-46710\T-47040 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46710, Common Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Dist_ATA_AT | bilateral | 125100\20168-1\G-A119\T-47040\T-47700 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Dist_Brachial_A_AT | bilateral | 125100\20168-1\G-A119\T-47020\T-47160 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47160, Brachial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Dist_CCA_AT | bilateral | 125100\20168-1\G-A119\T-45005\T-45100 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Dist_ICA_AT | bilateral | 125100\20168-1\G-A119\T-45005\T-45300 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Dist_Pero_A_AT | bilateral | 125100\20168-1\G-A119\T-47040\T-47630 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Dist_POPA_AT | bilateral | 125100\20168-1\G-A119\T-47040\T-47500 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47500, Popliteal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Dist_PTA_AT | bilateral | 125100\20168-1\G-A119\T-47040\T-47600 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Dist_Radial_A_AT | bilateral | 125100\20168-1\G-A119\T-47020\T-47300 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Dist_SFA_AT | bilateral | 125100\20168-1\G-A119\T-47040\T-47403 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Dist_Ulnar_A_AT | bilateral | 125100\20168-1\G-A119\T-47020\T-47200 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Dor_Ped_ATA_AT | bilateral | 125100\20168-1\T-47040\T-47741 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47741, Dorsalis Pedis Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Ext_Iliac_A_AT | bilateral | 125100\20168-1\T-46910\T-47040 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46910, External Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Int_Iliac_A_AT | bilateral | 125100\20168-1\T-46740\T-47040 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46740, Internal Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Mid_ATA_AT | bilateral | 125100\20168-1\G-A188\T-47040\T-47700 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Mid_CCA_AT | bilateral | 125100\20168-1\G-A188\T-45005\T-45100 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Mid_ECA_AT | bilateral | 125100\20168-1\G-A188\T-45005\T-45200 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-45200, External Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Mid_ICA_AT | bilateral | 125100\20168-1\G-A188\T-45005\T-45300 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Mid_Pero_A_AT | bilateral | 125100\20168-1\G-A188\T-47040\T-47630 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Mid_PTA_AT | bilateral | 125100\20168-1\G-A188\T-47040\T-47600 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Mid_Radial_A_AT | bilateral | 125100\20168-1\G-A188\T-47020\T-47300 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Mid_SCL_A_AT | bilateral | 125100\20168-1\G-A188\T-45005\T-46100 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46100, Subclavian Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Mid_SFA_AT | bilateral | 125100\20168-1\G-A188\T-47040\T-47403 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Mid_Ulnar_A_AT | bilateral | 125100\20168-1\G-A188\T-47020\T-47200 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_PFA_AT | bilateral | 125100\20168-1\G-A118\T-47040\T-47440 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47440, Profunda Femoris Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Prox_ATA_AT | bilateral | 125100\20168-1\G-A118\T-47040\T-47700 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Prox_Brachial_A_AT | bilateral | 125100\20168-1\G-A118\T-47020\T-47160 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47160, Brachial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Prox_CCA_AT | bilateral | 125100\20168-1\G-A118\T-45005\T-45100 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Prox_ECA_AT | bilateral | 125100\20168-1\G-A118\T-45005\T-45200 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-45200, External Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Prox_ICA_AT | bilateral | 125100\20168-1\G-A118\T-45005\T-45300 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Prox_Pero_A_AT | bilateral | 125100\20168-1\G-A118\T-47040\T-47630 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Prox_POPA_AT | bilateral | 125100\20168-1\G-A118\T-47040\T-47500 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47500, Popliteal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Prox_PTA_AT | bilateral | 125100\20168-1\G-A118\T-47040\T-47600 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Prox_Radial_A_AT | bilateral | 125100\20168-1\G-A118\T-47020\T-47300 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Prox_SCL_A_AT | bilateral | 125100\20168-1\G-A118\T-45005\T-46100 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46100, Subclavian Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Prox_SFA_AT | bilateral | 125100\20168-1\G-A118\T-47040\T-47403 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Prox_Ulnar_A_AT | bilateral | 125100\20168-1\G-A118\T-47020\T-47200 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | AT | Time | meas | m_Vertebral_A_AT | bilateral | 125100\20168-1\T-45005\T-45700 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-45700, Vertebral Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | Ax A PI | none | calc | c_Axillary_A_PI | bilateral | 125100\12008-9\T-47020\T-47100 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47100, Axillary Artery |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | Ax A RI | none | calc | c_Axillary_A_RI | bilateral | 125100\12023-8\T-47020\T-47100 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47100, Axillary Artery |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | Ax A S/D | none | calc | c_Axillary_A_SD | bilateral | 125100\12144-2\T-47020\T-47100 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47100, Axillary Artery |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | Bifur ACA EDV | Signed End Diastolic Vel | meas | m_tcd_Bi_ACA_EDV | bilateral | 125100\11653-3\G-0369\T-40501 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Bifur ACA PI | none | calc | c_tcd_Bi_ACA_PI | bilateral | 125100\12008-9\G-0369\T-40501 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Bifur ACA PSV | Signed Maximum Velocity | meas | m_tcd_Bi_ACA_PSV | bilateral | 125100\11726-7\G-0369\T-40501 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Bifur ACA TAPV | Signed Time Average Peak Vel | meas | m_tcd_Bi_ACA_TAPV | bilateral | 125100\11692-1\G-0369\T-40501 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Bifur ACA TCD Mean | none | calc | c_tcd_Bi_ACA_TcdMean | bilateral | 125100\C12221-03\G-0369\T-40501 |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Bifur ACA TCD PI | none | calc | c_tcd_Bi_ACA_TcdPI | bilateral | 125100\C12221-04\G-0369\T-40501 |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, G-0369, Anterior-Posterior Cerebral Artery Bifurcation |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Bifur MCA EDV | Signed End Diastolic Vel | meas | m_tcd_Bi_MCA_EDV | bilateral | 125100\11653-3\G-0368\T-40501 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Bifur MCA PI | none | calc | c_tcd_Bi_MCA_PI | bilateral | 125100\12008-9\G-0368\T-40501 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Bifur MCA PSV | Signed Maximum Velocity | meas | m_tcd_Bi_MCA_PSV | bilateral | 125100\11726-7\G-0368\T-40501 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Bifur MCA TAPV | Signed Time Average Peak Vel | meas | m_tcd_Bi_MCA_TAPV | bilateral | 125100\11692-1\G-0368\T-40501 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Bifur MCA TCD Mean | none | calc | c_tcd_Bi_MCA_TcdMean | bilateral | 125100\C12221-03\G-0368\T-40501 |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Bifur MCA TCD PI | none | calc | c_tcd_Bi_MCA_TcdPI | bilateral | 125100\C12221-04\G-0368\T-40501 |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, G-0368, Anterior-Middle Cerebral Artery Bifurcation |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Bladder Height | Length | meas | m_bladder_ht | unilateral | 125100\C12011-03\T5100-02\T-74000 |  | 99PMSBLUS, C12011-03, Bladder Height |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-74000, Bladder |  |  | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral)
Vascular Ultrasound Procedure Report | Vascular | Bladder Height | Length | meas | m_renal_bladder_ht | unilateral | 125100\C12011-03\T5100-02\T-74000 |  | 99PMSBLUS, C12011-03, Bladder Height |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-74000, Bladder |  |  | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral)
Vascular Ultrasound Procedure Report | Vascular | Bladder Length | Length | meas | m_bladder_len | unilateral | 125100\C12011-01\T5100-02\T-74000 |  | 99PMSBLUS, C12011-01, Bladder Length |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-74000, Bladder |  |  | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral)
Vascular Ultrasound Procedure Report | Vascular | Bladder Length | Length | meas | m_renal_bladder_len | unilateral | 125100\C12011-01\T5100-02\T-74000 |  | 99PMSBLUS, C12011-01, Bladder Length |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-74000, Bladder |  |  | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral)
Vascular Ultrasound Procedure Report | Vascular | Bladder Vol | none | calc | c_bladder_vol | unilateral | 125100\C12011-04\T5100-02\T-74000 |  | 99PMSBLUS, C12011-04, Bladder Volume |  | UCUM, cm3, Cubic Centimeter | [6 empty] | SRT, T-74000, Bladder |  |  | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral)
Vascular Ultrasound Procedure Report | Vascular | Bladder Vol | none | calc | c_renal_bladder_vol | unilateral | 125100\C12011-04\T5100-02\T-74000 |  | 99PMSBLUS, C12011-04, Bladder Volume |  | UCUM, cm3, Cubic Centimeter | [6 empty] | SRT, T-74000, Bladder |  |  | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral)
Vascular Ultrasound Procedure Report | Vascular | Bladder Width | Length | meas | m_bladder_wid | unilateral | 125100\C12011-02\T5100-02\T-74000 |  | 99PMSBLUS, C12011-02, Bladder Width |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-74000, Bladder |  |  | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral)
Vascular Ultrasound Procedure Report | Vascular | Bladder Width | Length | meas | m_renal_bladder_wid | unilateral | 125100\C12011-02\T5100-02\T-74000 |  | 99PMSBLUS, C12011-02, Bladder Width |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-74000, Bladder |  |  | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral)
Vascular Ultrasound Procedure Report | Vascular | Brachioceph A PI | none | calc | c_Brach_Ceph_A_PI | bilateral | 125100\12008-9\T-46010\T-47020 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46010, Innominate Artery |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | Brachioceph A RI | none | calc | c_Brach_Ceph_A_RI | bilateral | 125100\12023-8\T-46010\T-47020 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46010, Innominate Artery |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | Brachioceph A S/D | none | calc | c_Brach_Ceph_A_SD | bilateral | 125100\12144-2\T-46010\T-47020 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46010, Innominate Artery |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | Bulb IMT(Q) | Length | meas | m_Bulb_IMT | bilateral | 125100\C12122-08\C3627-01\T-45005\T-45170 |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45170, Carotid Bulb |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | Bulb PI | none | calc | c_Bulb_PI | bilateral | 125100\12008-9\T-45005\T-45170 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45170, Carotid Bulb |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | Bulb RI | none | calc | c_Bulb_RI | bilateral | 125100\12023-8\T-45005\T-45170 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45170, Carotid Bulb |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | Bulb S/D | none | calc | c_Bulb_SD | bilateral | 125100\12144-2\T-45005\T-45170 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-45170, Carotid Bulb |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | Carotid Siphon EDV | Signed End Diastolic Vel | meas | m_tcd_Car_Siph_EDV | bilateral | 125100\11653-3\T-40501\T-45308 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45308, Carotid Siphon |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Carotid Siphon PI | none | calc | c_tcd_Car_Siph_PI | bilateral | 125100\12008-9\T-40501\T-45308 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45308, Carotid Siphon |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Carotid Siphon PSV | Signed Maximum Velocity | meas | m_tcd_Car_Siph_PSV | bilateral | 125100\11726-7\T-40501\T-45308 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45308, Carotid Siphon |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Carotid Siphon TAPV | Signed Time Average Peak Vel | meas | m_tcd_Car_Siph_TAPV | bilateral | 125100\11692-1\T-40501\T-45308 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45308, Carotid Siphon |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Carotid Siphon TCD Mean | none | calc | c_tcd_Car_Siph_TcdMean | bilateral | 125100\C12221-03\T-40501\T-45308 |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45308, Carotid Siphon |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Carotid Siphon TCD PI | none | calc | c_tcd_Car_Siph_TcdPI | bilateral | 125100\C12221-04\T-40501\T-45308 |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45308, Carotid Siphon |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | CBD Diam | Length | meas | m_cbd | unilateral | 125100\121206\T5100-02\T-60610 |  | DCM, 121206, Distance |  | UCUM, mm, Millimeter | [6 empty] | SRT, T-60610, Bileduct |  |  | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral)
Vascular Ultrasound Procedure Report | Vascular | CCA % Area Reduction | none | calc | c_CCA_Area_Reduc | bilateral | 125100\G-0371\T-45005\T-45100 |  | SRT, G-0371, Percent Area Reduction |  | UCUM, %, Percent | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | CCA % Diam Reduction | none | calc | c_CCA_Diam_Reduc | bilateral | 125100\G-0372\T-45005\T-45100 |  | SRT, G-0372, Percent Diameter Reduction |  | UCUM, %, Percent | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | CCA Area Resid | Area | meas | m_CCA_Area_Resid | bilateral | 125100\C7471-01\T-45005\T-45100 |  | 99PMSBLUS, C7471-01, Area 1 of Area Percent Reduction |  | UCUM, cm2, Square Centimeter | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | CCA Area True | Area | meas | m_CCA_Area_True | bilateral | 125100\C7471-02\T-45005\T-45100 |  | 99PMSBLUS, C7471-02, Area 2 of Area Percent Reduction |  | UCUM, cm2, Square Centimeter | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | CCA Diam Resid | Length | meas | m_CCA_Diam_Resid | bilateral | 125100\C7470-05\T-45005\T-45100 |  | 99PMSBLUS, C7470-05, Diameter 1 of Diameter Reduction |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | CCA Diam True | Length | meas | m_CCA_Diam_True | bilateral | 125100\C7470-06\T-45005\T-45100 |  | 99PMSBLUS, C7470-06, Diameter 2 of Diameter Reduction |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | CCA IMT(Q) | Length | meas | m_CCA_IMT | bilateral | 125100\C12122-08\C3627-01\T-45005\T-45100 |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | CCA PSV | Signed Maximum Velocity | meas | m_CCA | bilateral | 125100\11726-7\C12104-02\T-45005 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | 99PMSBLUS, C12104-02, ICA/CCA Ratio Denominator |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | Celiac A AI | Doppler Slope | meas | m_celiacaxis_AS | unilateral | 125100\20167-3\T-46002\T-46400 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46400, Celiac Axis |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Celiac A AT | Time | meas | m_celiacaxis_AT | unilateral | 125100\20168-1\T-46002\T-46400 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46400, Celiac Axis |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Celiac A EDV | Signed End Diastolic Vel | meas | m_celiacaxis_EDV | unilateral | 125100\11653-3\T-46002\T-46400 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46400, Celiac Axis |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Celiac A MDV | Signed Minimum Velocity | meas | m_celiacaxis_MDV | unilateral | 125100\11665-7\T-46002\T-46400 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46400, Celiac Axis |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Celiac A PI | none | calc | c_celiacaxis_PI | unilateral | 125100\12008-9\T-46002\T-46400 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46400, Celiac Axis |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Celiac A PSV | Signed Maximum Velocity | meas | m_celiacaxis_PSV | unilateral | 125100\11726-7\T-46002\T-46400 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46400, Celiac Axis |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Celiac A RI | none | calc | c_celiacaxis_RI | unilateral | 125100\12023-8\T-46002\T-46400 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46400, Celiac Axis |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Celiac A S/D | none | calc | c_celiacaxis_SD | unilateral | 125100\12144-2\T-46002\T-46400 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46400, Celiac Axis |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Celiac A TAMV | Time Average Mean Velocity | meas | m_celiacaxis_TAMV | unilateral | 125100\20352-1\T-46002\T-46400 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46400, Celiac Axis |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Celiac A TAPV | Signed Time Average Peak Vel | meas | m_celiacaxis_TAPV | unilateral | 125100\11692-1\T-46002\T-46400 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46400, Celiac Axis |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Celiac A VTI | Velocity Time Integral | meas | m_celiacaxis_VTI | unilateral | 125100\20354-7\T-46002\T-46400 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46400, Celiac Axis |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | CFA % Area Reduction | none | calc | c_CFA_Area_Reduc | bilateral | 125100\G-0371\T-47040\T-47400 |  | SRT, G-0371, Percent Area Reduction |  | UCUM, %, Percent | [6 empty] | SRT, T-47400, Common Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | CFA % Diam Reduction | none | calc | c_CFA_Diam_Reduc | bilateral | 125100\G-0372\T-47040\T-47400 |  | SRT, G-0372, Percent Diameter Reduction |  | UCUM, %, Percent | [6 empty] | SRT, T-47400, Common Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | CFA Area Resid | Area | meas | m_CFA_Area_Resid | bilateral | 125100\C7471-01\T-47040\T-47400 |  | 99PMSBLUS, C7471-01, Area 1 of Area Percent Reduction |  | UCUM, cm2, Square Centimeter | [6 empty] | SRT, T-47400, Common Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | CFA Area True | Area | meas | m_CFA_Area_True | bilateral | 125100\C7471-02\T-47040\T-47400 |  | 99PMSBLUS, C7471-02, Area 2 of Area Percent Reduction |  | UCUM, cm2, Square Centimeter | [6 empty] | SRT, T-47400, Common Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | CFA Diam Resid | Length | meas | m_CFA_Diam_Resid | bilateral | 125100\C7470-05\T-47040\T-47400 |  | 99PMSBLUS, C7470-05, Diameter 1 of Diameter Reduction |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47400, Common Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | CFA Diam True | Length | meas | m_CFA_Diam_True | bilateral | 125100\C7470-06\T-47040\T-47400 |  | 99PMSBLUS, C7470-06, Diameter 2 of Diameter Reduction |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47400, Common Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | CFA PI | none | calc | c_CFA_PI | bilateral | 125100\12008-9\T-47040\T-47400 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47400, Common Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | CFA RI | none | calc | c_CFA_RI | bilateral | 125100\12023-8\T-47040\T-47400 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47400, Common Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | CFA S/D | none | calc | c_CFA_SD | bilateral | 125100\12144-2\T-47040\T-47400 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47400, Common Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | Com Hepatic A AI | Doppler Slope | meas | m_comhepart_AS | unilateral | 125100\20167-3\T-46002\T-46421 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46421, Common Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Com Hepatic A AT | Time | meas | m_comhepart_AT | unilateral | 125100\20168-1\T-46002\T-46421 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46421, Common Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Com Hepatic A EDV | Signed End Diastolic Vel | meas | m_comhepart_EDV | unilateral | 125100\11653-3\T-46002\T-46421 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46421, Common Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Com Hepatic A MDV | Signed Minimum Velocity | meas | m_comhepart_MDV | unilateral | 125100\11665-7\T-46002\T-46421 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46421, Common Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Com Hepatic A PI | none | calc | c_comhepart_PI | unilateral | 125100\12008-9\T-46002\T-46421 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46421, Common Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Com Hepatic A PSV | Signed Maximum Velocity | meas | m_comhepart_PSV | unilateral | 125100\11726-7\T-46002\T-46421 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46421, Common Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Com Hepatic A RI | none | calc | c_comhepart_RI | unilateral | 125100\12023-8\T-46002\T-46421 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46421, Common Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Com Hepatic A S/D | none | calc | c_comhepart_SD | unilateral | 125100\12144-2\T-46002\T-46421 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46421, Common Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Com Hepatic A TAMV | Time Average Mean Velocity | meas | m_comhepart_TAMV | unilateral | 125100\20352-1\T-46002\T-46421 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46421, Common Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Com Hepatic A TAPV | Signed Time Average Peak Vel | meas | m_comhepart_TAPV | unilateral | 125100\11692-1\T-46002\T-46421 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46421, Common Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Com Hepatic A VTI | Velocity Time Integral | meas | m_comhepart_VTI | unilateral | 125100\20354-7\T-46002\T-46421 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46421, Common Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Com Iliac A PI | none | calc | c_Com_Iliac_A_PI | bilateral | 125100\12008-9\T-46710\T-47040 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46710, Common Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | Com Iliac A RI | none | calc | c_Com_Iliac_A_RI | bilateral | 125100\12023-8\T-46710\T-47040 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46710, Common Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | Com Iliac A S/D | none | calc | c_Com_Iliac_A_SD | bilateral | 125100\12144-2\T-46710\T-47040 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46710, Common Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | Dist Ao AI | Doppler Slope | meas | m_distaorta_AS | unilateral | 125100\20167-3\G-A119\T-42000\T-46002 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-42000, Aorta |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ao AT | Time | meas | m_distaorta_AT | unilateral | 125100\20168-1\G-A119\T-42000\T-46002 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-42000, Aorta |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ao EDV | Signed End Diastolic Vel | meas | m_distaorta_EDV | unilateral | 125100\11653-3\G-A119\T-42000\T-46002 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-42000, Aorta |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ao MDV | Signed Minimum Velocity | meas | m_distaorta_MDV | unilateral | 125100\11665-7\G-A119\T-42000\T-46002 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-42000, Aorta |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ao PI | none | calc | c_distaorta_PI | unilateral | 125100\12008-9\G-A119\T-42000\T-46002 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-42000, Aorta |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ao PSV | Signed Maximum Velocity | meas | m_distaorta_PSV | unilateral | 125100\11726-7\G-A119\T-42000\T-46002 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-42000, Aorta |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ao RI | none | calc | c_distaorta_RI | unilateral | 125100\12023-8\G-A119\T-42000\T-46002 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-42000, Aorta |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ao S/D | none | calc | c_distaorta_SD | unilateral | 125100\12144-2\G-A119\T-42000\T-46002 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-42000, Aorta |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ao TAMV | Time Average Mean Velocity | meas | m_distaorta_TAMV | unilateral | 125100\20352-1\G-A119\T-42000\T-46002 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-42000, Aorta |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ao TAPV | Signed Time Average Peak Vel | meas | m_distaorta_TAPV | unilateral | 125100\11692-1\G-A119\T-42000\T-46002 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-42000, Aorta |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ao VTI | Velocity Time Integral | meas | m_distaorta_VTI | unilateral | 125100\20354-7\G-A119\T-42000\T-46002 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-42000, Aorta |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATA PI | none | calc | c_Dist_ATA_PI | bilateral | 125100\12008-9\G-A119\T-47040\T-47700 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATA RI | none | calc | c_Dist_ATA_RI | bilateral | 125100\12023-8\G-A119\T-47040\T-47700 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ATA S/D | none | calc | c_Dist_ATA_SD | bilateral | 125100\12144-2\G-A119\T-47040\T-47700 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilar A EDV | Signed End Diastolic Vel | meas | m_tcd_Dist_Bas_A_EDV | unilateral | 125100\11653-3\G-A119\T-40501\T-45800 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45800, Basilar Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilar A PI | none | calc | c_tcd_Dist_Bas_A_PI | unilateral | 125100\12008-9\G-A119\T-40501\T-45800 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45800, Basilar Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilar A PSV | Signed Maximum Velocity | meas | m_tcd_Dist_Bas_A_PSV | unilateral | 125100\11726-7\G-A119\T-40501\T-45800 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45800, Basilar Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilar A TAPV | Signed Time Average Peak Vel | meas | m_tcd_Dist_Bas_A_TAPV | unilateral | 125100\11692-1\G-A119\T-40501\T-45800 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45800, Basilar Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilar A TCD Mean | none | calc | c_tcd_Dist_Bas_A_TcdMean | unilateral | 125100\C12221-03\G-A119\T-40501\T-45800 |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45800, Basilar Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Basilar A TCD PI | none | calc | c_tcd_Dist_Bas_A_TcdPI | unilateral | 125100\C12221-04\G-A119\T-40501\T-45800 |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45800, Basilar Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach A PI | none | calc | c_Dist_Brachial_A_PI | bilateral | 125100\12008-9\G-A119\T-47020\T-47160 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47160, Brachial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach A RI | none | calc | c_Dist_Brachial_A_RI | bilateral | 125100\12023-8\G-A119\T-47020\T-47160 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47160, Brachial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Brach A S/D | none | calc | c_Dist_Brachial_A_SD | bilateral | 125100\12144-2\G-A119\T-47020\T-47160 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47160, Brachial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Bulb IMT | Intimal Medial Thickness | meas | MEAS_DBulb_IMT | bilateral | 125100\C12122-08\G-A119\T-45005\T-45170 |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45170, Carotid Bulb |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Bulb IMT(Q) | Length | meas | m_Dist_Bulb_IMT | bilateral | 125100\C12122-08\C3627-01\G-A119\T-45005\T-45170 |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45170, Carotid Bulb |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA IMT | Intimal Medial Thickness | meas | MEAS_DCCA_IMT | bilateral | 125100\C12122-08\G-A119\T-45005\T-45100 |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA IMT(Q) | Length | meas | m_Dist_CCA_IMT | bilateral | 125100\C12122-08\C3627-01\G-A119\T-45005\T-45100 |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA PI | none | calc | c_Dist_CCA_PI | bilateral | 125100\12008-9\G-A119\T-45005\T-45100 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA RI | none | calc | c_Dist_CCA_RI | bilateral | 125100\12023-8\G-A119\T-45005\T-45100 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist CCA S/D | none | calc | c_Dist_CCA_SD | bilateral | 125100\12144-2\G-A119\T-45005\T-45100 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA EDV | Signed End Diastolic Vel | meas | m_tcd_Dist_ICA_EDV | bilateral | 125100\11653-3\G-A119\T-40501\T-45300 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA IMT | Intimal Medial Thickness | meas | MEAS_DICA_IMT | bilateral | 125100\C12122-08\G-A119\T-45005\T-45300 |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA IMT(Q) | Length | meas | m_Dist_ICA_IMT | bilateral | 125100\C12122-08\C3627-01\G-A119\T-45005\T-45300 |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA PI | none | calc | c_tcd_Dist_ICA_PI | bilateral | 125100\12008-9\G-A119\T-40501\T-45300 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA PI | none | calc | c_Dist_ICA_PI | bilateral | 125100\12008-9\G-A119\T-45005\T-45300 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA PSV | Signed Maximum Velocity | meas | m_tcd_Dist_ICA_PSV | bilateral | 125100\11726-7\G-A119\T-40501\T-45300 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA RI | none | calc | c_Dist_ICA_RI | bilateral | 125100\12023-8\G-A119\T-45005\T-45300 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA S/D | none | calc | c_Dist_ICA_SD | bilateral | 125100\12144-2\G-A119\T-45005\T-45300 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA TAPV | Signed Time Average Peak Vel | meas | m_tcd_Dist_ICA_TAPV | bilateral | 125100\11692-1\G-A119\T-40501\T-45300 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA TCD Mean | none | calc | c_tcd_Dist_ICA_TcdMean | bilateral | 125100\C12221-03\G-A119\T-40501\T-45300 |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist ICA TCD PI | none | calc | c_tcd_Dist_ICA_TcdPI | bilateral | 125100\C12221-04\G-A119\T-40501\T-45300 |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist M1 EDV | Signed End Diastolic Vel | meas | m_tcd_Dist_MCA_M1_EDV | bilateral | 125100\11653-3\G-A119\R-1024F\T-40501 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist M1 PI | none | calc | c_tcd_Dist_MCA_M1_PI | bilateral | 125100\12008-9\G-A119\R-1024F\T-40501 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist M1 PSV | Signed Maximum Velocity | meas | m_tcd_Dist_MCA_M1_PSV | bilateral | 125100\11726-7\G-A119\R-1024F\T-40501 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist M1 TAPV | Signed Time Average Peak Vel | meas | m_tcd_Dist_MCA_M1_TAPV | bilateral | 125100\11692-1\G-A119\R-1024F\T-40501 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist M1 TCD Mean | none | calc | c_tcd_Dist_MCA_M1_TcdMean | bilateral | 125100\C12221-03\G-A119\R-1024F\T-40501 |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist M1 TCD PI | none | calc | c_tcd_Dist_MCA_M1_TcdPI | bilateral | 125100\C12221-04\G-A119\R-1024F\T-40501 |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero A PI | none | calc | c_Dist_Pero_A_PI | bilateral | 125100\12008-9\G-A119\T-47040\T-47630 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero A RI | none | calc | c_Dist_Pero_A_RI | bilateral | 125100\12023-8\G-A119\T-47040\T-47630 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pero A S/D | none | calc | c_Dist_Pero_A_SD | bilateral | 125100\12144-2\G-A119\T-47040\T-47630 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop A PI | none | calc | c_Dist_POPA_PI | bilateral | 125100\12008-9\G-A119\T-47040\T-47500 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47500, Popliteal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop A RI | none | calc | c_Dist_POPA_RI | bilateral | 125100\12023-8\G-A119\T-47040\T-47500 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47500, Popliteal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Pop A S/D | none | calc | c_Dist_POPA_SD | bilateral | 125100\12144-2\G-A119\T-47040\T-47500 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47500, Popliteal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTA PI | none | calc | c_Dist_PTA_PI | bilateral | 125100\12008-9\G-A119\T-47040\T-47600 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTA RI | none | calc | c_Dist_PTA_RI | bilateral | 125100\12023-8\G-A119\T-47040\T-47600 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist PTA S/D | none | calc | c_Dist_PTA_SD | bilateral | 125100\12144-2\G-A119\T-47040\T-47600 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad A PI | none | calc | c_Dist_Radial_A_PI | bilateral | 125100\12008-9\G-A119\T-47020\T-47300 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad A RI | none | calc | c_Dist_Radial_A_RI | bilateral | 125100\12023-8\G-A119\T-47020\T-47300 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Rad A S/D | none | calc | c_Dist_Radial_A_SD | bilateral | 125100\12144-2\G-A119\T-47020\T-47300 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ren A AI | Doppler Slope | meas | m_distrenart_AS | bilateral | 125100\20167-3\G-A119\T-46600\T-71019 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ren A AT | Time | meas | m_distrenart_AT | bilateral | 125100\20168-1\G-A119\T-46600\T-71019 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ren A EDV | Signed End Diastolic Vel | meas | m_distrenart_EDV | bilateral | 125100\11653-3\G-A119\T-46600\T-71019 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ren A MDV | Signed Minimum Velocity | meas | m_distrenart_MDV | bilateral | 125100\11665-7\G-A119\T-46600\T-71019 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ren A PI | none | calc | c_distrenart_PI | bilateral | 125100\12008-9\G-A119\T-46600\T-71019 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ren A PSV | Signed Maximum Velocity | meas | m_distrenart_PSV | bilateral | 125100\11726-7\G-A119\T-46600\T-71019 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ren A RI | none | calc | c_distrenart_RI | bilateral | 125100\12023-8\G-A119\T-46600\T-71019 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ren A S/D | none | calc | c_distrenart_SD | bilateral | 125100\12144-2\G-A119\T-46600\T-71019 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ren A TAMV | Time Average Mean Velocity | meas | m_distrenart_TAMV | bilateral | 125100\20352-1\G-A119\T-46600\T-71019 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ren A TAPV | Signed Time Average Peak Vel | meas | m_distrenart_TAPV | bilateral | 125100\11692-1\G-A119\T-46600\T-71019 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ren A VTI | Velocity Time Integral | meas | m_distrenart_VTI | bilateral | 125100\20354-7\G-A119\T-46600\T-71019 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFA PI | none | calc | c_Dist_SFA_PI | bilateral | 125100\12008-9\G-A119\T-47040\T-47403 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFA RI | none | calc | c_Dist_SFA_RI | bilateral | 125100\12023-8\G-A119\T-47040\T-47403 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SFA S/D | none | calc | c_Dist_SFA_SD | bilateral | 125100\12144-2\G-A119\T-47040\T-47403 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SMA AI | Doppler Slope | meas | m_distsma_AS | unilateral | 125100\20167-3\G-A119\T-46002\T-46510 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SMA AT | Time | meas | m_distsma_AT | unilateral | 125100\20168-1\G-A119\T-46002\T-46510 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SMA EDV | Signed End Diastolic Vel | meas | m_distsma_EDV | unilateral | 125100\11653-3\G-A119\T-46002\T-46510 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SMA MDV | Signed Minimum Velocity | meas | m_distsma_MDV | unilateral | 125100\11665-7\G-A119\T-46002\T-46510 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SMA PI | none | calc | c_distsma_PI | unilateral | 125100\12008-9\G-A119\T-46002\T-46510 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SMA PSV | Signed Maximum Velocity | meas | m_distsma_PSV | unilateral | 125100\11726-7\G-A119\T-46002\T-46510 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SMA RI | none | calc | c_distsma_RI | unilateral | 125100\12023-8\G-A119\T-46002\T-46510 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SMA S/D | none | calc | c_distsma_SD | unilateral | 125100\12144-2\G-A119\T-46002\T-46510 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SMA TAMV | Time Average Mean Velocity | meas | m_distsma_TAMV | unilateral | 125100\20352-1\G-A119\T-46002\T-46510 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SMA TAPV | Signed Time Average Peak Vel | meas | m_distsma_TAPV | unilateral | 125100\11692-1\G-A119\T-46002\T-46510 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist SMA VTI | Velocity Time Integral | meas | m_distsma_VTI | unilateral | 125100\20354-7\G-A119\T-46002\T-46510 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar A PI | none | calc | c_Dist_Ulnar_A_PI | bilateral | 125100\12008-9\G-A119\T-47020\T-47200 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar A RI | none | calc | c_Dist_Ulnar_A_RI | bilateral | 125100\12023-8\G-A119\T-47020\T-47200 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dist Ulnar A S/D | none | calc | c_Dist_Ulnar_A_SD | bilateral | 125100\12144-2\G-A119\T-47020\T-47200 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | Dor Pedis PI | none | calc | c_Dor_Ped_ATA_PI | bilateral | 125100\12008-9\T-47040\T-47741 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47741, Dorsalis Pedis Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | Dor Pedis RI | none | calc | c_Dor_Ped_ATA_RI | bilateral | 125100\12023-8\T-47040\T-47741 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47741, Dorsalis Pedis Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | Dor Pedis S/D | none | calc | c_Dor_Ped_ATA_SD | bilateral | 125100\12144-2\T-47040\T-47741 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47741, Dorsalis Pedis Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Antecube_EDV | bilateral | 125100\11653-3\sup71_001\T-47020 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | 99PMSBLUS, sup71_001, Antecube |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Axillary_A_EDV | bilateral | 125100\11653-3\T-47020\T-47100 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47100, Axillary Artery |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Brach_Ceph_A_EDV | bilateral | 125100\11653-3\T-46010\T-47020 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46010, Innominate Artery |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Bulb_EDV | bilateral | 125100\11653-3\T-45005\T-45170 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45170, Carotid Bulb |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_CFA_EDV | bilateral | 125100\11653-3\T-47040\T-47400 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47400, Common Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Com_Iliac_A_EDV | bilateral | 125100\11653-3\T-46710\T-47040 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46710, Common Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Dist_ATA_EDV | bilateral | 125100\11653-3\G-A119\T-47040\T-47700 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Dist_Brachial_A_EDV | bilateral | 125100\11653-3\G-A119\T-47020\T-47160 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47160, Brachial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Dist_CCA_EDV | bilateral | 125100\11653-3\G-A119\T-45005\T-45100 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Dist_ICA_EDV | bilateral | 125100\11653-3\G-A119\T-45005\T-45300 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Dist_Pero_A_EDV | bilateral | 125100\11653-3\G-A119\T-47040\T-47630 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Dist_POPA_EDV | bilateral | 125100\11653-3\G-A119\T-47040\T-47500 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47500, Popliteal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Dist_PTA_EDV | bilateral | 125100\11653-3\G-A119\T-47040\T-47600 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Dist_Radial_A_EDV | bilateral | 125100\11653-3\G-A119\T-47020\T-47300 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Dist_SFA_EDV | bilateral | 125100\11653-3\G-A119\T-47040\T-47403 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Dist_Ulnar_A_EDV | bilateral | 125100\11653-3\G-A119\T-47020\T-47200 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Dor_Ped_ATA_EDV | bilateral | 125100\11653-3\T-47040\T-47741 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47741, Dorsalis Pedis Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Ext_Iliac_A_EDV | bilateral | 125100\11653-3\T-46910\T-47040 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46910, External Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Int_Iliac_A_EDV | bilateral | 125100\11653-3\T-46740\T-47040 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46740, Internal Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Mid_ATA_EDV | bilateral | 125100\11653-3\G-A188\T-47040\T-47700 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Mid_CCA_EDV | bilateral | 125100\11653-3\G-A188\T-45005\T-45100 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Mid_ECA_EDV | bilateral | 125100\11653-3\G-A188\T-45005\T-45200 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45200, External Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Mid_ICA_EDV | bilateral | 125100\11653-3\G-A188\T-45005\T-45300 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Mid_Pero_A_EDV | bilateral | 125100\11653-3\G-A188\T-47040\T-47630 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Mid_PTA_EDV | bilateral | 125100\11653-3\G-A188\T-47040\T-47600 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Mid_Radial_A_EDV | bilateral | 125100\11653-3\G-A188\T-47020\T-47300 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Mid_SCL_A_EDV | bilateral | 125100\11653-3\G-A188\T-45005\T-46100 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46100, Subclavian Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Mid_SFA_EDV | bilateral | 125100\11653-3\G-A188\T-47040\T-47403 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Mid_Ulnar_A_EDV | bilateral | 125100\11653-3\G-A188\T-47020\T-47200 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_PFA_EDV | bilateral | 125100\11653-3\G-A118\T-47040\T-47440 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47440, Profunda Femoris Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Prox_ATA_EDV | bilateral | 125100\11653-3\G-A118\T-47040\T-47700 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Prox_Brachial_A_EDV | bilateral | 125100\11653-3\G-A118\T-47020\T-47160 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47160, Brachial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Prox_CCA_EDV | bilateral | 125100\11653-3\G-A118\T-45005\T-45100 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Prox_ECA_EDV | bilateral | 125100\11653-3\G-A118\T-45005\T-45200 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45200, External Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Prox_ICA_EDV | bilateral | 125100\11653-3\G-A118\T-45005\T-45300 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Prox_Pero_A_EDV | bilateral | 125100\11653-3\G-A118\T-47040\T-47630 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Prox_POPA_EDV | bilateral | 125100\11653-3\G-A118\T-47040\T-47500 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47500, Popliteal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Prox_PTA_EDV | bilateral | 125100\11653-3\G-A118\T-47040\T-47600 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Prox_Radial_A_EDV | bilateral | 125100\11653-3\G-A118\T-47020\T-47300 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Prox_SCL_A_EDV | bilateral | 125100\11653-3\G-A118\T-45005\T-46100 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46100, Subclavian Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Prox_SFA_EDV | bilateral | 125100\11653-3\G-A118\T-47040\T-47403 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Prox_Ulnar_A_EDV | bilateral | 125100\11653-3\G-A118\T-47020\T-47200 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | EDV | Signed End Diastolic Vel | meas | m_Vertebral_A_EDV | bilateral | 125100\11653-3\T-45005\T-45700 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45700, Vertebral Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac A PI | none | calc | c_Ext_Iliac_A_PI | bilateral | 125100\12008-9\T-46910\T-47040 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46910, External Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac A RI | none | calc | c_Ext_Iliac_A_RI | bilateral | 125100\12023-8\T-46910\T-47040 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46910, External Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | Ext Iliac A S/D | none | calc | c_Ext_Iliac_A_SD | bilateral | 125100\12144-2\T-46910\T-47040 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46910, External Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | GB Wall Diam | Length | meas | m_gbWall | unilateral | 125100\121206\T5100-02\T-63000 |  | DCM, 121206, Distance |  | UCUM, mm, Millimeter | [6 empty] | SRT, T-63000, Gallbladder |  |  | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral)
Vascular Ultrasound Procedure Report | Vascular | GDA AI | Doppler Slope | meas | m_gda_AS | unilateral | 125100\20167-3\T-46002\T-46440 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46440, Gastroduodenal Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | GDA AT | Time | meas | m_gda_AT | unilateral | 125100\20168-1\T-46002\T-46440 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46440, Gastroduodenal Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | GDA EDV | Signed End Diastolic Vel | meas | m_gda_EDV | unilateral | 125100\11653-3\T-46002\T-46440 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46440, Gastroduodenal Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | GDA MDV | Signed Minimum Velocity | meas | m_gda_MDV | unilateral | 125100\11665-7\T-46002\T-46440 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46440, Gastroduodenal Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | GDA PI | none | calc | c_gda_PI | unilateral | 125100\12008-9\T-46002\T-46440 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46440, Gastroduodenal Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | GDA PSV | Signed Maximum Velocity | meas | m_gda_PSV | unilateral | 125100\11726-7\T-46002\T-46440 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46440, Gastroduodenal Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | GDA RI | none | calc | c_gda_RI | unilateral | 125100\12023-8\T-46002\T-46440 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46440, Gastroduodenal Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | GDA S/D | none | calc | c_gda_SD | unilateral | 125100\12144-2\T-46002\T-46440 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46440, Gastroduodenal Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | GDA TAMV | Time Average Mean Velocity | meas | m_gda_TAMV | unilateral | 125100\20352-1\T-46002\T-46440 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46440, Gastroduodenal Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | GDA TAPV | Signed Time Average Peak Vel | meas | m_gda_TAPV | unilateral | 125100\11692-1\T-46002\T-46440 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46440, Gastroduodenal Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | GDA VTI | Velocity Time Integral | meas | m_gda_VTI | unilateral | 125100\20354-7\T-46002\T-46440 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46440, Gastroduodenal Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Hilar A AI | Doppler Slope | meas | m_hilarart_AS | bilateral | 125100\20167-3\G-035C\T-71019 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, G-035C, Hilar Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Hilar A AT | Time | meas | m_hilarart_AT | bilateral | 125100\20168-1\G-035C\T-71019 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, G-035C, Hilar Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Hilar A EDV | Signed End Diastolic Vel | meas | m_hilarart_EDV | bilateral | 125100\11653-3\G-035C\T-71019 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, G-035C, Hilar Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Hilar A MDV | Signed Minimum Velocity | meas | m_hilarart_MDV | bilateral | 125100\11665-7\G-035C\T-71019 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, G-035C, Hilar Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Hilar A PI | none | calc | c_hilarart_PI | bilateral | 125100\12008-9\G-035C\T-71019 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, G-035C, Hilar Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Hilar A PSV | Signed Maximum Velocity | meas | m_hilarart_PSV | bilateral | 125100\11726-7\G-035C\T-71019 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, G-035C, Hilar Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Hilar A RI | none | calc | c_hilarart_RI | bilateral | 125100\12023-8\G-035C\T-71019 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, G-035C, Hilar Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Hilar A S/D | none | calc | c_hilarart_SD | bilateral | 125100\12144-2\G-035C\T-71019 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, G-035C, Hilar Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Hilar A TAMV | Time Average Mean Velocity | meas | m_hilarart_TAMV | bilateral | 125100\20352-1\G-035C\T-71019 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, G-035C, Hilar Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Hilar A TAPV | Signed Time Average Peak Vel | meas | m_hilarart_TAPV | bilateral | 125100\11692-1\G-035C\T-71019 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, G-035C, Hilar Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Hilar A VTI | Velocity Time Integral | meas | m_hilarart_VTI | bilateral | 125100\20354-7\G-035C\T-71019 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, G-035C, Hilar Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | ICA | Unsigned Maximum Velocity | meas | m_tcd_ICA | bilateral | 125100\C12105-02\C12221-03\T-40501 |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | 99PMSBLUS, C12105-02, MCA/Dist ICA Ratio Denominator |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | ICA % Area Reduction | none | calc | c_ICA_Area_Reduc | bilateral | 125100\G-0371\T-45005\T-45300 |  | SRT, G-0371, Percent Area Reduction |  | UCUM, %, Percent | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | ICA % Diam Reduction | none | calc | c_ICA_Diam_Reduc | bilateral | 125100\G-0372\T-45005\T-45300 |  | SRT, G-0372, Percent Diameter Reduction |  | UCUM, %, Percent | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | ICA Area Resid | Area | meas | m_ICA_Area_Resid | bilateral | 125100\C7471-01\T-45005\T-45300 |  | 99PMSBLUS, C7471-01, Area 1 of Area Percent Reduction |  | UCUM, cm2, Square Centimeter | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | ICA Area True | Area | meas | m_ICA_Area_True | bilateral | 125100\C7471-02\T-45005\T-45300 |  | 99PMSBLUS, C7471-02, Area 2 of Area Percent Reduction |  | UCUM, cm2, Square Centimeter | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | ICA Diam Resid | Length | meas | m_ICA_Diam_Resid | bilateral | 125100\C7470-05\T-45005\T-45300 |  | 99PMSBLUS, C7470-05, Diameter 1 of Diameter Reduction |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | ICA Diam True | Length | meas | m_ICA_Diam_True | bilateral | 125100\C7470-06\T-45005\T-45300 |  | 99PMSBLUS, C7470-06, Diameter 2 of Diameter Reduction |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | ICA IMT(Q) | Length | meas | m_ICA_IMT | bilateral | 125100\C12122-08\C3627-01\T-45005\T-45300 |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | ICA PSV | Signed Maximum Velocity | meas | m_ICA | bilateral | 125100\11726-7\C12104-01\T-45005 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | 99PMSBLUS, C12104-01, ICA/CCA Ratio Numerator |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | ICA/CCA | none | calc | c_ICACCA | bilateral | 125100\33868-1\T-45005 |  | LN, 33868-1, ICA/CCA velocity ratio |  | UCUM, 1, ratio | [9 empty] | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | IMA AI | Doppler Slope | meas | m_ima_AS | unilateral | 125100\20167-3\T-46002\T-46520 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46520, Inferior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | IMA AT | Time | meas | m_ima_AT | unilateral | 125100\20168-1\T-46002\T-46520 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46520, Inferior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | IMA EDV | Signed End Diastolic Vel | meas | m_ima_EDV | unilateral | 125100\11653-3\T-46002\T-46520 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46520, Inferior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | IMA MDV | Signed Minimum Velocity | meas | m_ima_MDV | unilateral | 125100\11665-7\T-46002\T-46520 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46520, Inferior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | IMA PI | none | calc | c_ima_PI | unilateral | 125100\12008-9\T-46002\T-46520 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46520, Inferior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | IMA PSV | Signed Maximum Velocity | meas | m_ima_PSV | unilateral | 125100\11726-7\T-46002\T-46520 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46520, Inferior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | IMA RI | none | calc | c_ima_RI | unilateral | 125100\12023-8\T-46002\T-46520 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46520, Inferior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | IMA S/D | none | calc | c_ima_SD | unilateral | 125100\12144-2\T-46002\T-46520 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46520, Inferior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | IMA TAMV | Time Average Mean Velocity | meas | m_ima_TAMV | unilateral | 125100\20352-1\T-46002\T-46520 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46520, Inferior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | IMA TAPV | Signed Time Average Peak Vel | meas | m_ima_TAPV | unilateral | 125100\11692-1\T-46002\T-46520 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46520, Inferior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | IMA VTI | Velocity Time Integral | meas | m_ima_VTI | unilateral | 125100\20354-7\T-46002\T-46520 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46520, Inferior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Inf Arc AI | Doppler Slope | meas | m_infarc_AS | bilateral | 125100\20167-3\G-A115\T-4668A\T-71019 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Inf Arc AT | Time | meas | m_infarc_AT | bilateral | 125100\20168-1\G-A115\T-4668A\T-71019 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Inf Arc EDV | Signed End Diastolic Vel | meas | m_infarc_EDV | bilateral | 125100\11653-3\G-A115\T-4668A\T-71019 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Inf Arc MDV | Signed Minimum Velocity | meas | m_infarc_MDV | bilateral | 125100\11665-7\G-A115\T-4668A\T-71019 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Inf Arc PI | none | calc | c_infarc_PI | bilateral | 125100\12008-9\G-A115\T-4668A\T-71019 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Inf Arc PSV | Signed Maximum Velocity | meas | m_infarc_PSV | bilateral | 125100\11726-7\G-A115\T-4668A\T-71019 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Inf Arc RI | none | calc | c_infarc_RI | bilateral | 125100\12023-8\G-A115\T-4668A\T-71019 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Inf Arc S/D | none | calc | c_infarc_SD | bilateral | 125100\12144-2\G-A115\T-4668A\T-71019 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Inf Arc TAMV | Time Average Mean Velocity | meas | m_infarc_TAMV | bilateral | 125100\20352-1\G-A115\T-4668A\T-71019 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Inf Arc TAPV | Signed Time Average Peak Vel | meas | m_infarc_TAPV | bilateral | 125100\11692-1\G-A115\T-4668A\T-71019 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Inf Arc VTI | Velocity Time Integral | meas | m_infarc_VTI | bilateral | 125100\20354-7\G-A115\T-4668A\T-71019 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A115, Inferior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Inf Ren Ao AI | Doppler Slope | meas | m_infrenaorta_AS | unilateral | 125100\20167-3\T-42520\T-46002 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-42520, Infra-renal Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Inf Ren Ao AT | Time | meas | m_infrenaorta_AT | unilateral | 125100\20168-1\T-42520\T-46002 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-42520, Infra-renal Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Inf Ren Ao EDV | Signed End Diastolic Vel | meas | m_infrenaorta_EDV | unilateral | 125100\11653-3\T-42520\T-46002 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-42520, Infra-renal Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Inf Ren Ao MDV | Signed Minimum Velocity | meas | m_infrenaorta_MDV | unilateral | 125100\11665-7\T-42520\T-46002 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-42520, Infra-renal Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Inf Ren Ao PI | none | calc | c_infrenaorta_PI | unilateral | 125100\12008-9\T-42520\T-46002 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-42520, Infra-renal Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Inf Ren Ao PSV | Signed Maximum Velocity | meas | m_infrenaorta_PSV | unilateral | 125100\11726-7\T-42520\T-46002 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-42520, Infra-renal Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Inf Ren Ao RI | none | calc | c_infrenaorta_RI | unilateral | 125100\12023-8\T-42520\T-46002 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-42520, Infra-renal Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Inf Ren Ao S/D | none | calc | c_infrenaorta_SD | unilateral | 125100\12144-2\T-42520\T-46002 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-42520, Infra-renal Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Inf Ren Ao TAMV | Time Average Mean Velocity | meas | m_infrenaorta_TAMV | unilateral | 125100\20352-1\T-42520\T-46002 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-42520, Infra-renal Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Inf Ren Ao TAPV | Signed Time Average Peak Vel | meas | m_infrenaorta_TAPV | unilateral | 125100\11692-1\T-42520\T-46002 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-42520, Infra-renal Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Inf Ren Ao VTI | Velocity Time Integral | meas | m_infrenaorta_VTI | unilateral | 125100\20354-7\T-42520\T-46002 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-42520, Infra-renal Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Inf Seg AI | Doppler Slope | meas | m_infseg_AS | bilateral | 125100\20167-3\G-A115\T-46659\T-71019 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Inf Seg AT | Time | meas | m_infseg_AT | bilateral | 125100\20168-1\G-A115\T-46659\T-71019 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Inf Seg EDV | Signed End Diastolic Vel | meas | m_infseg_EDV | bilateral | 125100\11653-3\G-A115\T-46659\T-71019 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Inf Seg MDV | Signed Minimum Velocity | meas | m_infseg_MDV | bilateral | 125100\11665-7\G-A115\T-46659\T-71019 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Inf Seg PI | none | calc | c_infseg_PI | bilateral | 125100\12008-9\G-A115\T-46659\T-71019 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Inf Seg PSV | Signed Maximum Velocity | meas | m_infseg_PSV | bilateral | 125100\11726-7\G-A115\T-46659\T-71019 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Inf Seg RI | none | calc | c_infseg_RI | bilateral | 125100\12023-8\G-A115\T-46659\T-71019 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Inf Seg S/D | none | calc | c_infseg_SD | bilateral | 125100\12144-2\G-A115\T-46659\T-71019 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Inf Seg TAMV | Time Average Mean Velocity | meas | m_infseg_TAMV | bilateral | 125100\20352-1\G-A115\T-46659\T-71019 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Inf Seg TAPV | Signed Time Average Peak Vel | meas | m_infseg_TAPV | bilateral | 125100\11692-1\G-A115\T-46659\T-71019 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Inf Seg VTI | Velocity Time Integral | meas | m_infseg_VTI | bilateral | 125100\20354-7\G-A115\T-46659\T-71019 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A115, Inferior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Int Iliac A PI | none | calc | c_Int_Iliac_A_PI | bilateral | 125100\12008-9\T-46740\T-47040 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46740, Internal Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | Int Iliac A RI | none | calc | c_Int_Iliac_A_RI | bilateral | 125100\12023-8\T-46740\T-47040 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46740, Internal Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | Int Iliac A S/D | none | calc | c_Int_Iliac_A_SD | bilateral | 125100\12144-2\T-46740\T-47040 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46740, Internal Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | IVC Diameter | Length | meas | m_IVCDiam_Abd | unilateral | 125100\18006-7\T-48710\T-487A0 |  | LN, 18006-7, Inferior Vena Cava Diameter |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-48710, Inferior Vena Cava |  |  | SRT, T-487A0, Vein of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Kidney Height | Length | meas | m_general_kidht | bilateral | 125100\C7470-01\T5100-01\T-71000 |  | 99PMSBLUS, C7470-01, Height |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-71000, Kidney |  |  | 99PMSBLUS, T5100-01, Anatomic Structures
Vascular Ultrasound Procedure Report | Vascular | Kidney Height | Length | meas | m_kidht | bilateral | 125100\C7470-01\T5100-01\T-71000 |  | 99PMSBLUS, C7470-01, Height |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-71000, Kidney |  |  | 99PMSBLUS, T5100-01, Anatomic Structures
Vascular Ultrasound Procedure Report | Vascular | Kidney Length | Length | meas | m_general_kidlen | bilateral | 125100\G-A22A\T5100-01\T-71000 |  | SRT, G-A22A, Length |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-71000, Kidney |  |  | 99PMSBLUS, T5100-01, Anatomic Structures
Vascular Ultrasound Procedure Report | Vascular | Kidney Length | Length | meas | m_kidlen | bilateral | 125100\G-A22A\T5100-01\T-71000 |  | SRT, G-A22A, Length |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-71000, Kidney |  |  | 99PMSBLUS, T5100-01, Anatomic Structures
Vascular Ultrasound Procedure Report | Vascular | Kidney Vol | none | calc | c_general_kidvol | bilateral | 125100\121221\T5100-01\T-71000 |  | DCM, 121221, Volume of ellipsoid |  | UCUM, cm3, Cubic Centimeter | [6 empty] | SRT, T-71000, Kidney |  |  | 99PMSBLUS, T5100-01, Anatomic Structures
Vascular Ultrasound Procedure Report | Vascular | Kidney Vol | none | calc | c_kidvol | bilateral | 125100\121221\T5100-01\T-71000 |  | DCM, 121221, Volume of ellipsoid |  | UCUM, cm3, Cubic Centimeter | [6 empty] | SRT, T-71000, Kidney |  |  | 99PMSBLUS, T5100-01, Anatomic Structures
Vascular Ultrasound Procedure Report | Vascular | Kidney Width | Length | meas | m_general_kidwidth | bilateral | 125100\G-A220\T5100-01\T-71000 |  | SRT, G-A220, Width |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-71000, Kidney |  |  | 99PMSBLUS, T5100-01, Anatomic Structures
Vascular Ultrasound Procedure Report | Vascular | Kidney Width | Length | meas | m_kidwidth | bilateral | 125100\G-A220\T5100-01\T-71000 |  | SRT, G-A220, Width |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-71000, Kidney |  |  | 99PMSBLUS, T5100-01, Anatomic Structures
Vascular Ultrasound Procedure Report | Vascular | L Hepatic A AI | Doppler Slope | meas | m_l_hepart_AS | unilateral | 125100\20167-3\T-46002\T-46427 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46427, Left Branch of Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | L Hepatic A AT | Time | meas | m_l_hepart_AT | unilateral | 125100\20168-1\T-46002\T-46427 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46427, Left Branch of Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | L Hepatic A EDV | Signed End Diastolic Vel | meas | m_l_hepart_EDV | unilateral | 125100\11653-3\T-46002\T-46427 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46427, Left Branch of Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | L Hepatic A MDV | Signed Minimum Velocity | meas | m_l_hepart_MDV | unilateral | 125100\11665-7\T-46002\T-46427 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46427, Left Branch of Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | L Hepatic A PI | none | calc | c_l_hepart_PI | unilateral | 125100\12008-9\T-46002\T-46427 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46427, Left Branch of Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | L Hepatic A PSV | Signed Maximum Velocity | meas | m_l_hepart_PSV | unilateral | 125100\11726-7\T-46002\T-46427 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46427, Left Branch of Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | L Hepatic A RI | none | calc | c_l_hepart_RI | unilateral | 125100\12023-8\T-46002\T-46427 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46427, Left Branch of Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | L Hepatic A S/D | none | calc | c_l_hepart_SD | unilateral | 125100\12144-2\T-46002\T-46427 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46427, Left Branch of Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | L Hepatic A TAMV | Time Average Mean Velocity | meas | m_l_hepart_TAMV | unilateral | 125100\20352-1\T-46002\T-46427 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46427, Left Branch of Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | L Hepatic A TAPV | Signed Time Average Peak Vel | meas | m_l_hepart_TAPV | unilateral | 125100\11692-1\T-46002\T-46427 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46427, Left Branch of Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | L Hepatic A VTI | Velocity Time Integral | meas | m_l_hepart_VTI | unilateral | 125100\20354-7\T-46002\T-46427 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46427, Left Branch of Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Liver Length | Length | meas | m_liverLen | unilateral | 125100\G-A22A\T5100-02\T-62000 |  | SRT, G-A22A, Length |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-62000, Liver |  |  | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral)
Vascular Ultrasound Procedure Report | Vascular | Liver Vasc Angle | Alpha Angle | meas | m_liver_vasc_angle | unilateral | 125100\C12221-05\T5100-02\T-62000 |  | 99PMSBLUS, C12221-05, Liver Vasc Angle |  | UCUM, deg, deg | [6 empty] | SRT, T-62000, Liver |  |  | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral)
Vascular Ultrasound Procedure Report | Vascular | Lung Angle | Alpha Angle | meas | m_lung_angle | bilateral | 125100\C7470-22\T-28000\T5100-01 |  | 99PMSBLUS, C7470-22, Angle |  | UCUM, deg, deg | [6 empty] | SRT, T-28000, Lung |  |  | 99PMSBLUS, T5100-01, Anatomic Structures
Vascular Ultrasound Procedure Report | Vascular | M2 EDV | Signed End Diastolic Vel | meas | m_tcd_MCA_M2_EDV | bilateral | 125100\11653-3\R-10251\T-40501 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | M2 PI | none | calc | c_tcd_MCA_M2_PI | bilateral | 125100\12008-9\R-10251\T-40501 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | M2 PSV | Signed Maximum Velocity | meas | m_tcd_MCA_M2_PSV | bilateral | 125100\11726-7\R-10251\T-40501 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | M2 TAPV | Signed Time Average Peak Vel | meas | m_tcd_MCA_M2_TAPV | bilateral | 125100\11692-1\R-10251\T-40501 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | M2 TCD Mean | none | calc | c_tcd_MCA_M2_TcdMean | bilateral | 125100\C12221-03\R-10251\T-40501 |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | M2 TCD PI | none | calc | c_tcd_MCA_M2_TcdPI | bilateral | 125100\C12221-04\R-10251\T-40501 |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, R-10251, Middle Cerebral Artery M2 Segment |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | MCA | Unsigned Maximum Velocity | meas | m_tcd_MCA | bilateral | 125100\C12105-01\C12221-03\T-40501 |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | 99PMSBLUS, C12105-01, MCA/Dist ICA Ratio Numerator |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | MCA/ICA | none | calc | c_tcd_MCAICA | bilateral | 125100\C12105-03\C12105-05\T-40501 |  | 99PMSBLUS, C12105-03, MCA/Dist ICA Ratio |  | UCUM, 1, ratio | [6 empty] | 99PMSBLUS, C12105-05, MCA/Dist ICA Ratio |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Antecube_MDV | bilateral | 125100\11665-7\sup71_001\T-47020 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | 99PMSBLUS, sup71_001, Antecube |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Axillary_A_MDV | bilateral | 125100\11665-7\T-47020\T-47100 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47100, Axillary Artery |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Brach_Ceph_A_MDV | bilateral | 125100\11665-7\T-46010\T-47020 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46010, Innominate Artery |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Bulb_MDV | bilateral | 125100\11665-7\T-45005\T-45170 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45170, Carotid Bulb |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_CFA_MDV | bilateral | 125100\11665-7\T-47040\T-47400 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47400, Common Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Com_Iliac_A_MDV | bilateral | 125100\11665-7\T-46710\T-47040 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46710, Common Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Dist_ATA_MDV | bilateral | 125100\11665-7\G-A119\T-47040\T-47700 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Dist_Brachial_A_MDV | bilateral | 125100\11665-7\G-A119\T-47020\T-47160 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47160, Brachial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Dist_CCA_MDV | bilateral | 125100\11665-7\G-A119\T-45005\T-45100 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Dist_ICA_MDV | bilateral | 125100\11665-7\G-A119\T-45005\T-45300 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Dist_Pero_A_MDV | bilateral | 125100\11665-7\G-A119\T-47040\T-47630 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Dist_POPA_MDV | bilateral | 125100\11665-7\G-A119\T-47040\T-47500 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47500, Popliteal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Dist_PTA_MDV | bilateral | 125100\11665-7\G-A119\T-47040\T-47600 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Dist_Radial_A_MDV | bilateral | 125100\11665-7\G-A119\T-47020\T-47300 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Dist_SFA_MDV | bilateral | 125100\11665-7\G-A119\T-47040\T-47403 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Dist_Ulnar_A_MDV | bilateral | 125100\11665-7\G-A119\T-47020\T-47200 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Dor_Ped_ATA_MDV | bilateral | 125100\11665-7\T-47040\T-47741 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47741, Dorsalis Pedis Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Ext_Iliac_A_MDV | bilateral | 125100\11665-7\T-46910\T-47040 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46910, External Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Int_Iliac_A_MDV | bilateral | 125100\11665-7\T-46740\T-47040 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46740, Internal Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Mid_ATA_MDV | bilateral | 125100\11665-7\G-A188\T-47040\T-47700 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Mid_CCA_MDV | bilateral | 125100\11665-7\G-A188\T-45005\T-45100 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Mid_ECA_MDV | bilateral | 125100\11665-7\G-A188\T-45005\T-45200 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45200, External Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Mid_ICA_MDV | bilateral | 125100\11665-7\G-A188\T-45005\T-45300 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Mid_Pero_A_MDV | bilateral | 125100\11665-7\G-A188\T-47040\T-47630 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Mid_PTA_MDV | bilateral | 125100\11665-7\G-A188\T-47040\T-47600 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Mid_Radial_A_MDV | bilateral | 125100\11665-7\G-A188\T-47020\T-47300 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Mid_SCL_A_MDV | bilateral | 125100\11665-7\G-A188\T-45005\T-46100 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46100, Subclavian Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Mid_SFA_MDV | bilateral | 125100\11665-7\G-A188\T-47040\T-47403 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Mid_Ulnar_A_MDV | bilateral | 125100\11665-7\G-A188\T-47020\T-47200 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_PFA_MDV | bilateral | 125100\11665-7\G-A118\T-47040\T-47440 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47440, Profunda Femoris Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Prox_ATA_MDV | bilateral | 125100\11665-7\G-A118\T-47040\T-47700 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Prox_Brachial_A_MDV | bilateral | 125100\11665-7\G-A118\T-47020\T-47160 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47160, Brachial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Prox_CCA_MDV | bilateral | 125100\11665-7\G-A118\T-45005\T-45100 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Prox_ECA_MDV | bilateral | 125100\11665-7\G-A118\T-45005\T-45200 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45200, External Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Prox_ICA_MDV | bilateral | 125100\11665-7\G-A118\T-45005\T-45300 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Prox_Pero_A_MDV | bilateral | 125100\11665-7\G-A118\T-47040\T-47630 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Prox_POPA_MDV | bilateral | 125100\11665-7\G-A118\T-47040\T-47500 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47500, Popliteal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Prox_PTA_MDV | bilateral | 125100\11665-7\G-A118\T-47040\T-47600 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Prox_Radial_A_MDV | bilateral | 125100\11665-7\G-A118\T-47020\T-47300 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Prox_SCL_A_MDV | bilateral | 125100\11665-7\G-A118\T-45005\T-46100 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46100, Subclavian Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Prox_SFA_MDV | bilateral | 125100\11665-7\G-A118\T-47040\T-47403 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Prox_Ulnar_A_MDV | bilateral | 125100\11665-7\G-A118\T-47020\T-47200 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | MDV | Signed Minimum Velocity | meas | m_Vertebral_A_MDV | bilateral | 125100\11665-7\T-45005\T-45700 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45700, Vertebral Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | Mid Arc AI | Doppler Slope | meas | m_medarc_AS | bilateral | 125100\20167-3\G-A109\T-4668A\T-71019 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Mid Arc AT | Time | meas | m_medarc_AT | bilateral | 125100\20168-1\G-A109\T-4668A\T-71019 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Mid Arc EDV | Signed End Diastolic Vel | meas | m_medarc_EDV | bilateral | 125100\11653-3\G-A109\T-4668A\T-71019 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Mid Arc MDV | Signed Minimum Velocity | meas | m_medarc_MDV | bilateral | 125100\11665-7\G-A109\T-4668A\T-71019 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Mid Arc PI | none | calc | c_medarc_PI | bilateral | 125100\12008-9\G-A109\T-4668A\T-71019 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Mid Arc PSV | Signed Maximum Velocity | meas | m_medarc_PSV | bilateral | 125100\11726-7\G-A109\T-4668A\T-71019 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Mid Arc RI | none | calc | c_medarc_RI | bilateral | 125100\12023-8\G-A109\T-4668A\T-71019 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Mid Arc S/D | none | calc | c_medarc_SD | bilateral | 125100\12144-2\G-A109\T-4668A\T-71019 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Mid Arc TAMV | Time Average Mean Velocity | meas | m_medarc_TAMV | bilateral | 125100\20352-1\G-A109\T-4668A\T-71019 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Mid Arc TAPV | Signed Time Average Peak Vel | meas | m_medarc_TAPV | bilateral | 125100\11692-1\G-A109\T-4668A\T-71019 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Mid Arc VTI | Velocity Time Integral | meas | m_medarc_VTI | bilateral | 125100\20354-7\G-A109\T-4668A\T-71019 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A109, Medial |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Mid ATA PI | none | calc | c_Mid_ATA_PI | bilateral | 125100\12008-9\G-A188\T-47040\T-47700 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATA RI | none | calc | c_Mid_ATA_RI | bilateral | 125100\12023-8\G-A188\T-47040\T-47700 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ATA S/D | none | calc | c_Mid_ATA_SD | bilateral | 125100\12144-2\G-A188\T-47040\T-47700 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilar A EDV | Signed End Diastolic Vel | meas | m_tcd_Mid_Bas_A_EDV | unilateral | 125100\11653-3\G-A188\T-40501\T-45800 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45800, Basilar Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilar A PI | none | calc | c_tcd_Mid_Bas_A_PI | unilateral | 125100\12008-9\G-A188\T-40501\T-45800 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45800, Basilar Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilar A PSV | Signed Maximum Velocity | meas | m_tcd_Mid_Bas_A_PSV | unilateral | 125100\11726-7\G-A188\T-40501\T-45800 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45800, Basilar Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilar A TAPV | Signed Time Average Peak Vel | meas | m_tcd_Mid_Bas_A_TAPV | unilateral | 125100\11692-1\G-A188\T-40501\T-45800 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45800, Basilar Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilar A TCD Mean | none | calc | c_tcd_Mid_Bas_A_TcdMean | unilateral | 125100\C12221-03\G-A188\T-40501\T-45800 |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45800, Basilar Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Basilar A TCD PI | none | calc | c_tcd_Mid_Bas_A_TcdPI | unilateral | 125100\C12221-04\G-A188\T-40501\T-45800 |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45800, Basilar Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Bulb IMT | Intimal Medial Thickness | meas | MEAS_MBulb_IMT | bilateral | 125100\C12122-08\G-A188\T-45005\T-45170 |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45170, Carotid Bulb |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Bulb IMT(Q) | Length | meas | m_Mid_Bulb_IMT | bilateral | 125100\C12122-08\C3627-01\G-A188\T-45005\T-45170 |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45170, Carotid Bulb |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA IMT | Intimal Medial Thickness | meas | MEAS_MCCA_IMT | bilateral | 125100\C12122-08\G-A188\T-45005\T-45100 |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA IMT(Q) | Length | meas | m_Mid_CCA_IMT | bilateral | 125100\C12122-08\C3627-01\G-A188\T-45005\T-45100 |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA PI | none | calc | c_Mid_CCA_PI | bilateral | 125100\12008-9\G-A188\T-45005\T-45100 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA RI | none | calc | c_Mid_CCA_RI | bilateral | 125100\12023-8\G-A188\T-45005\T-45100 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid CCA S/D | none | calc | c_Mid_CCA_SD | bilateral | 125100\12144-2\G-A188\T-45005\T-45100 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ECA PI | none | calc | c_Mid_ECA_PI | bilateral | 125100\12008-9\G-A188\T-45005\T-45200 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45200, External Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ECA RI | none | calc | c_Mid_ECA_RI | bilateral | 125100\12023-8\G-A188\T-45005\T-45200 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45200, External Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ECA S/D | none | calc | c_Mid_ECA_SD | bilateral | 125100\12144-2\G-A188\T-45005\T-45200 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-45200, External Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA IMT | Intimal Medial Thickness | meas | MEAS_MICA_IMT | bilateral | 125100\C12122-08\G-A188\T-45005\T-45300 |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA IMT(Q) | Length | meas | m_Mid_ICA_IMT | bilateral | 125100\C12122-08\C3627-01\G-A188\T-45005\T-45300 |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA PI | none | calc | c_Mid_ICA_PI | bilateral | 125100\12008-9\G-A188\T-45005\T-45300 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA RI | none | calc | c_Mid_ICA_RI | bilateral | 125100\12023-8\G-A188\T-45005\T-45300 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid ICA S/D | none | calc | c_Mid_ICA_SD | bilateral | 125100\12144-2\G-A188\T-45005\T-45300 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid M1 EDV | Signed End Diastolic Vel | meas | m_tcd_Mid_MCA_M1_EDV | bilateral | 125100\11653-3\G-A188\R-1024F\T-40501 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid M1 PI | none | calc | c_tcd_Mid_MCA_M1_PI | bilateral | 125100\12008-9\G-A188\R-1024F\T-40501 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid M1 PSV | Signed Maximum Velocity | meas | m_tcd_Mid_MCA_M1_PSV | bilateral | 125100\11726-7\G-A188\R-1024F\T-40501 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid M1 TAPV | Signed Time Average Peak Vel | meas | m_tcd_Mid_MCA_M1_TAPV | bilateral | 125100\11692-1\G-A188\R-1024F\T-40501 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid M1 TCD Mean | none | calc | c_tcd_Mid_MCA_M1_TcdMean | bilateral | 125100\C12221-03\G-A188\R-1024F\T-40501 |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid M1 TCD PI | none | calc | c_tcd_Mid_MCA_M1_TcdPI | bilateral | 125100\C12221-04\G-A188\R-1024F\T-40501 |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero A PI | none | calc | c_Mid_Pero_A_PI | bilateral | 125100\12008-9\G-A188\T-47040\T-47630 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero A RI | none | calc | c_Mid_Pero_A_RI | bilateral | 125100\12023-8\G-A188\T-47040\T-47630 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Pero A S/D | none | calc | c_Mid_Pero_A_SD | bilateral | 125100\12144-2\G-A188\T-47040\T-47630 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTA PI | none | calc | c_Mid_PTA_PI | bilateral | 125100\12008-9\G-A188\T-47040\T-47600 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTA RI | none | calc | c_Mid_PTA_RI | bilateral | 125100\12023-8\G-A188\T-47040\T-47600 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid PTA S/D | none | calc | c_Mid_PTA_SD | bilateral | 125100\12144-2\G-A188\T-47040\T-47600 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad A PI | none | calc | c_Mid_Radial_A_PI | bilateral | 125100\12008-9\G-A188\T-47020\T-47300 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad A RI | none | calc | c_Mid_Radial_A_RI | bilateral | 125100\12023-8\G-A188\T-47020\T-47300 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Rad A S/D | none | calc | c_Mid_Radial_A_SD | bilateral | 125100\12144-2\G-A188\T-47020\T-47300 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ren A AI | Doppler Slope | meas | m_midrenart_AS | bilateral | 125100\20167-3\G-A188\T-46600\T-71019 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ren A AT | Time | meas | m_midrenart_AT | bilateral | 125100\20168-1\G-A188\T-46600\T-71019 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ren A EDV | Signed End Diastolic Vel | meas | m_midrenart_EDV | bilateral | 125100\11653-3\G-A188\T-46600\T-71019 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ren A MDV | Signed Minimum Velocity | meas | m_midrenart_MDV | bilateral | 125100\11665-7\G-A188\T-46600\T-71019 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ren A PI | none | calc | c_midrenart_PI | bilateral | 125100\12008-9\G-A188\T-46600\T-71019 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ren A PSV | Signed Maximum Velocity | meas | m_midrenart_PSV | bilateral | 125100\11726-7\G-A188\T-46600\T-71019 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ren A RI | none | calc | c_midrenart_RI | bilateral | 125100\12023-8\G-A188\T-46600\T-71019 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ren A S/D | none | calc | c_midrenart_SD | bilateral | 125100\12144-2\G-A188\T-46600\T-71019 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ren A TAMV | Time Average Mean Velocity | meas | m_midrenart_TAMV | bilateral | 125100\20352-1\G-A188\T-46600\T-71019 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ren A TAPV | Signed Time Average Peak Vel | meas | m_midrenart_TAPV | bilateral | 125100\11692-1\G-A188\T-46600\T-71019 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ren A VTI | Velocity Time Integral | meas | m_midrenart_VTI | bilateral | 125100\20354-7\G-A188\T-46600\T-71019 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL A PI | none | calc | c_Mid_SCL_A_PI | bilateral | 125100\12008-9\G-A188\T-45005\T-46100 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46100, Subclavian Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL A RI | none | calc | c_Mid_SCL_A_RI | bilateral | 125100\12023-8\G-A188\T-45005\T-46100 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46100, Subclavian Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SCL A S/D | none | calc | c_Mid_SCL_A_SD | bilateral | 125100\12144-2\G-A188\T-45005\T-46100 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46100, Subclavian Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Seg AI | Doppler Slope | meas | m_medseg_AS | bilateral | 125100\20167-3\G-A109\T-46659\T-71019 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Mid Seg AT | Time | meas | m_medseg_AT | bilateral | 125100\20168-1\G-A109\T-46659\T-71019 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Mid Seg EDV | Signed End Diastolic Vel | meas | m_medseg_EDV | bilateral | 125100\11653-3\G-A109\T-46659\T-71019 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Mid Seg MDV | Signed Minimum Velocity | meas | m_medseg_MDV | bilateral | 125100\11665-7\G-A109\T-46659\T-71019 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Mid Seg PI | none | calc | c_medseg_PI | bilateral | 125100\12008-9\G-A109\T-46659\T-71019 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Mid Seg PSV | Signed Maximum Velocity | meas | m_medseg_PSV | bilateral | 125100\11726-7\G-A109\T-46659\T-71019 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Mid Seg RI | none | calc | c_medseg_RI | bilateral | 125100\12023-8\G-A109\T-46659\T-71019 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Mid Seg S/D | none | calc | c_medseg_SD | bilateral | 125100\12144-2\G-A109\T-46659\T-71019 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Mid Seg TAMV | Time Average Mean Velocity | meas | m_medseg_TAMV | bilateral | 125100\20352-1\G-A109\T-46659\T-71019 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Mid Seg TAPV | Signed Time Average Peak Vel | meas | m_medseg_TAPV | bilateral | 125100\11692-1\G-A109\T-46659\T-71019 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Mid Seg VTI | Velocity Time Integral | meas | m_medseg_VTI | bilateral | 125100\20354-7\G-A109\T-46659\T-71019 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A109, Medial |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Mid SFA PI | none | calc | c_Mid_SFA_PI | bilateral | 125100\12008-9\G-A188\T-47040\T-47403 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFA RI | none | calc | c_Mid_SFA_RI | bilateral | 125100\12023-8\G-A188\T-47040\T-47403 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SFA S/D | none | calc | c_Mid_SFA_SD | bilateral | 125100\12144-2\G-A188\T-47040\T-47403 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SMA AI | Doppler Slope | meas | m_midsma_AS | unilateral | 125100\20167-3\G-A188\T-46002\T-46510 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SMA AT | Time | meas | m_midsma_AT | unilateral | 125100\20168-1\G-A188\T-46002\T-46510 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SMA EDV | Signed End Diastolic Vel | meas | m_midsma_EDV | unilateral | 125100\11653-3\G-A188\T-46002\T-46510 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SMA MDV | Signed Minimum Velocity | meas | m_midsma_MDV | unilateral | 125100\11665-7\G-A188\T-46002\T-46510 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SMA PI | none | calc | c_midsma_PI | unilateral | 125100\12008-9\G-A188\T-46002\T-46510 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SMA PSV | Signed Maximum Velocity | meas | m_midsma_PSV | unilateral | 125100\11726-7\G-A188\T-46002\T-46510 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SMA RI | none | calc | c_midsma_RI | unilateral | 125100\12023-8\G-A188\T-46002\T-46510 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SMA S/D | none | calc | c_midsma_SD | unilateral | 125100\12144-2\G-A188\T-46002\T-46510 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SMA TAMV | Time Average Mean Velocity | meas | m_midsma_TAMV | unilateral | 125100\20352-1\G-A188\T-46002\T-46510 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SMA TAPV | Signed Time Average Peak Vel | meas | m_midsma_TAPV | unilateral | 125100\11692-1\G-A188\T-46002\T-46510 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid SMA VTI | Velocity Time Integral | meas | m_midsma_VTI | unilateral | 125100\20354-7\G-A188\T-46002\T-46510 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar A PI | none | calc | c_Mid_Ulnar_A_PI | bilateral | 125100\12008-9\G-A188\T-47020\T-47200 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar A RI | none | calc | c_Mid_Ulnar_A_RI | bilateral | 125100\12023-8\G-A188\T-47020\T-47200 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Mid Ulnar A S/D | none | calc | c_Mid_Ulnar_A_SD | bilateral | 125100\12144-2\G-A188\T-47020\T-47200 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | Ophthalmic A EDV | Signed End Diastolic Vel | meas | m_tcd_Oph_A_EDV | bilateral | 125100\11653-3\T-40501\T-45400 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45400, Ophthalmic Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Ophthalmic A PI | none | calc | c_tcd_Oph_A_PI | bilateral | 125100\12008-9\T-40501\T-45400 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45400, Ophthalmic Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Ophthalmic A PSV | Signed Maximum Velocity | meas | m_tcd_Oph_A_PSV | bilateral | 125100\11726-7\T-40501\T-45400 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45400, Ophthalmic Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Ophthalmic A TAPV | Signed Time Average Peak Vel | meas | m_tcd_Oph_A_TAPV | bilateral | 125100\11692-1\T-40501\T-45400 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45400, Ophthalmic Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Ophthalmic A TCD Mean | none | calc | c_tcd_Oph_A_TcdMean | bilateral | 125100\C12221-03\T-40501\T-45400 |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45400, Ophthalmic Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Ophthalmic A TCD PI | none | calc | c_tcd_Oph_A_TcdPI | bilateral | 125100\C12221-04\T-40501\T-45400 |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45400, Ophthalmic Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | PCA P1 EDV | Signed End Diastolic Vel | meas | m_tcd_PCA_P1_EDV | bilateral | 125100\11653-3\R-10253\T-40501 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | PCA P1 PI | none | calc | c_tcd_PCA_P1_PI | bilateral | 125100\12008-9\R-10253\T-40501 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | PCA P1 PSV | Signed Maximum Velocity | meas | m_tcd_PCA_P1_PSV | bilateral | 125100\11726-7\R-10253\T-40501 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | PCA P1 TAPV | Signed Time Average Peak Vel | meas | m_tcd_PCA_P1_TAPV | bilateral | 125100\11692-1\R-10253\T-40501 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | PCA P1 TCD Mean | none | calc | c_tcd_PCA_P1_TcdMean | bilateral | 125100\C12221-03\R-10253\T-40501 |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | PCA P1 TCD PI | none | calc | c_tcd_PCA_P1_TcdPI | bilateral | 125100\C12221-04\R-10253\T-40501 |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, R-10253, Posterior Cerebral Artery P1 Segment |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | PCA P2 EDV | Signed End Diastolic Vel | meas | m_tcd_PCA_P2_EDV | bilateral | 125100\11653-3\R-10255\T-40501 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | PCA P2 PI | none | calc | c_tcd_PCA_P2_PI | bilateral | 125100\12008-9\R-10255\T-40501 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | PCA P2 PSV | Signed Maximum Velocity | meas | m_tcd_PCA_P2_PSV | bilateral | 125100\11726-7\R-10255\T-40501 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | PCA P2 TAPV | Signed Time Average Peak Vel | meas | m_tcd_PCA_P2_TAPV | bilateral | 125100\11692-1\R-10255\T-40501 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | PCA P2 TCD Mean | none | calc | c_tcd_PCA_P2_TcdMean | bilateral | 125100\C12221-03\R-10255\T-40501 |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | PCA P2 TCD PI | none | calc | c_tcd_PCA_P2_TcdPI | bilateral | 125100\C12221-04\R-10255\T-40501 |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, R-10255, Posterior Cerebral Artery P2 Segment |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | PCoA EDV | Signed End Diastolic Vel | meas | m_tcd_PComA_EDV | bilateral | 125100\11653-3\T-40501\T-45320 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45320, Posterior Communicating Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | PCoA PI | none | calc | c_tcd_PComA_PI | bilateral | 125100\12008-9\T-40501\T-45320 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45320, Posterior Communicating Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | PCoA PSV | Signed Maximum Velocity | meas | m_tcd_PComA_PSV | bilateral | 125100\11726-7\T-40501\T-45320 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45320, Posterior Communicating Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | PCoA TAPV | Signed Time Average Peak Vel | meas | m_tcd_PComA_TAPV | bilateral | 125100\11692-1\T-40501\T-45320 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45320, Posterior Communicating Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | PCoA TCD Mean | none | calc | c_tcd_PComA_TcdMean | bilateral | 125100\C12221-03\T-40501\T-45320 |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45320, Posterior Communicating Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | PCoA TCD PI | none | calc | c_tcd_PComA_TcdPI | bilateral | 125100\C12221-04\T-40501\T-45320 |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45320, Posterior Communicating Artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Prox ATA PI | none | calc | c_Prox_ATA_PI | bilateral | 125100\12008-9\G-A118\T-47040\T-47700 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATA RI | none | calc | c_Prox_ATA_RI | bilateral | 125100\12023-8\G-A118\T-47040\T-47700 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ATA S/D | none | calc | c_Prox_ATA_SD | bilateral | 125100\12144-2\G-A118\T-47040\T-47700 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Basilar A EDV | Signed End Diastolic Vel | meas | m_tcd_Prox_Bas_A_EDV | unilateral | 125100\11653-3\G-A118\T-40501\T-45800 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45800, Basilar Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Basilar A PI | none | calc | c_tcd_Prox_Bas_A_PI | unilateral | 125100\12008-9\G-A118\T-40501\T-45800 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45800, Basilar Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Basilar A PSV | Signed Maximum Velocity | meas | m_tcd_Prox_Bas_A_PSV | unilateral | 125100\11726-7\G-A118\T-40501\T-45800 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45800, Basilar Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Basilar A TAPV | Signed Time Average Peak Vel | meas | m_tcd_Prox_Bas_A_TAPV | unilateral | 125100\11692-1\G-A118\T-40501\T-45800 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45800, Basilar Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Basilar A TCD Mean | none | calc | c_tcd_Prox_Bas_A_TcdMean | unilateral | 125100\C12221-03\G-A118\T-40501\T-45800 |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45800, Basilar Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Basilar A TCD PI | none | calc | c_tcd_Prox_Bas_A_TcdPI | unilateral | 125100\C12221-04\G-A118\T-40501\T-45800 |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45800, Basilar Artery |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach A PI | none | calc | c_Prox_Brachial_A_PI | bilateral | 125100\12008-9\G-A118\T-47020\T-47160 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47160, Brachial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach A RI | none | calc | c_Prox_Brachial_A_RI | bilateral | 125100\12023-8\G-A118\T-47020\T-47160 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47160, Brachial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Brach A S/D | none | calc | c_Prox_Brachial_A_SD | bilateral | 125100\12144-2\G-A118\T-47020\T-47160 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47160, Brachial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Bulb IMT | Intimal Medial Thickness | meas | MEAS_PBulb_IMT | bilateral | 125100\C12122-08\G-A118\T-45005\T-45170 |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45170, Carotid Bulb |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Bulb IMT(Q) | Length | meas | m_Prox_Bulb_IMT | bilateral | 125100\C12122-08\C3627-01\G-A118\T-45005\T-45170 |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45170, Carotid Bulb |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA IMT | Intimal Medial Thickness | meas | MEAS_PCCA_IMT | bilateral | 125100\C12122-08\G-A118\T-45005\T-45100 |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA IMT(Q) | Length | meas | m_Prox_CCA_IMT | bilateral | 125100\C12122-08\C3627-01\G-A118\T-45005\T-45100 |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA PI | none | calc | c_Prox_CCA_PI | bilateral | 125100\12008-9\G-A118\T-45005\T-45100 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA RI | none | calc | c_Prox_CCA_RI | bilateral | 125100\12023-8\G-A118\T-45005\T-45100 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox CCA S/D | none | calc | c_Prox_CCA_SD | bilateral | 125100\12144-2\G-A118\T-45005\T-45100 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ECA PI | none | calc | c_Prox_ECA_PI | bilateral | 125100\12008-9\G-A118\T-45005\T-45200 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45200, External Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ECA RI | none | calc | c_Prox_ECA_RI | bilateral | 125100\12023-8\G-A118\T-45005\T-45200 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45200, External Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ECA S/D | none | calc | c_Prox_ECA_SD | bilateral | 125100\12144-2\G-A118\T-45005\T-45200 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-45200, External Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA IMT | Intimal Medial Thickness | meas | MEAS_PICA_IMT | bilateral | 125100\C12122-08\G-A118\T-45005\T-45300 |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA IMT(Q) | Length | meas | m_Prox_ICA_IMT | bilateral | 125100\C12122-08\C3627-01\G-A118\T-45005\T-45300 |  | 99PMSBLUS, C12122-08, Mean Intima Media Thickness |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA PI | none | calc | c_Prox_ICA_PI | bilateral | 125100\12008-9\G-A118\T-45005\T-45300 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA RI | none | calc | c_Prox_ICA_RI | bilateral | 125100\12023-8\G-A118\T-45005\T-45300 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox ICA S/D | none | calc | c_Prox_ICA_SD | bilateral | 125100\12144-2\G-A118\T-45005\T-45300 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox M1 EDV | Signed End Diastolic Vel | meas | m_tcd_Prox_MCA_M1_EDV | bilateral | 125100\11653-3\G-A118\R-1024F\T-40501 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox M1 PI | none | calc | c_tcd_Prox_MCA_M1_PI | bilateral | 125100\12008-9\G-A118\R-1024F\T-40501 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox M1 PSV | Signed Maximum Velocity | meas | m_tcd_Prox_MCA_M1_PSV | bilateral | 125100\11726-7\G-A118\R-1024F\T-40501 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox M1 TAPV | Signed Time Average Peak Vel | meas | m_tcd_Prox_MCA_M1_TAPV | bilateral | 125100\11692-1\G-A118\R-1024F\T-40501 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox M1 TCD Mean | none | calc | c_tcd_Prox_MCA_M1_TcdMean | bilateral | 125100\C12221-03\G-A118\R-1024F\T-40501 |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox M1 TCD PI | none | calc | c_tcd_Prox_MCA_M1_TcdPI | bilateral | 125100\C12221-04\G-A118\R-1024F\T-40501 |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, R-1024F, Middle Cerebral Artery M1 Segment |  |  | SRT, T-40501, Blood Vessel of Head | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero A PI | none | calc | c_Prox_Pero_A_PI | bilateral | 125100\12008-9\G-A118\T-47040\T-47630 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero A RI | none | calc | c_Prox_Pero_A_RI | bilateral | 125100\12023-8\G-A118\T-47040\T-47630 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pero A S/D | none | calc | c_Prox_Pero_A_SD | bilateral | 125100\12144-2\G-A118\T-47040\T-47630 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PFA PI | none | calc | c_PFA_PI | bilateral | 125100\12008-9\G-A118\T-47040\T-47440 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47440, Profunda Femoris Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PFA RI | none | calc | c_PFA_RI | bilateral | 125100\12023-8\G-A118\T-47040\T-47440 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47440, Profunda Femoris Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PFA S/D | none | calc | c_PFA_SD | bilateral | 125100\12144-2\G-A118\T-47040\T-47440 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47440, Profunda Femoris Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop A PI | none | calc | c_Prox_POPA_PI | bilateral | 125100\12008-9\G-A118\T-47040\T-47500 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47500, Popliteal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop A RI | none | calc | c_Prox_POPA_RI | bilateral | 125100\12023-8\G-A118\T-47040\T-47500 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47500, Popliteal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Pop A S/D | none | calc | c_Prox_POPA_SD | bilateral | 125100\12144-2\G-A118\T-47040\T-47500 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47500, Popliteal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTA PI | none | calc | c_Prox_PTA_PI | bilateral | 125100\12008-9\G-A118\T-47040\T-47600 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTA RI | none | calc | c_Prox_PTA_RI | bilateral | 125100\12023-8\G-A118\T-47040\T-47600 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox PTA S/D | none | calc | c_Prox_PTA_SD | bilateral | 125100\12144-2\G-A118\T-47040\T-47600 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad A PI | none | calc | c_Prox_Radial_A_PI | bilateral | 125100\12008-9\G-A118\T-47020\T-47300 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad A RI | none | calc | c_Prox_Radial_A_RI | bilateral | 125100\12023-8\G-A118\T-47020\T-47300 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Rad A S/D | none | calc | c_Prox_Radial_A_SD | bilateral | 125100\12144-2\G-A118\T-47020\T-47300 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ren A AI | Doppler Slope | meas | m_proxrenart_AS | bilateral | 125100\20167-3\G-A118\T-46600\T-71019 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ren A AT | Time | meas | m_proxrenart_AT | bilateral | 125100\20168-1\G-A118\T-46600\T-71019 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ren A EDV | Signed End Diastolic Vel | meas | m_proxrenart_EDV | bilateral | 125100\11653-3\G-A118\T-46600\T-71019 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ren A MDV | Signed Minimum Velocity | meas | m_proxrenart_MDV | bilateral | 125100\11665-7\G-A118\T-46600\T-71019 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ren A PI | none | calc | c_proxrenart_PI | bilateral | 125100\12008-9\G-A118\T-46600\T-71019 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ren A PSV | Signed Maximum Velocity | meas | m_proxrenart_PSV | bilateral | 125100\11726-7\G-A118\T-46600\T-71019 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ren A RI | none | calc | c_proxrenart_RI | bilateral | 125100\12023-8\G-A118\T-46600\T-71019 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ren A S/D | none | calc | c_proxrenart_SD | bilateral | 125100\12144-2\G-A118\T-46600\T-71019 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ren A TAMV | Time Average Mean Velocity | meas | m_proxrenart_TAMV | bilateral | 125100\20352-1\G-A118\T-46600\T-71019 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ren A TAPV | Signed Time Average Peak Vel | meas | m_proxrenart_TAPV | bilateral | 125100\11692-1\G-A118\T-46600\T-71019 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ren A VTI | Velocity Time Integral | meas | m_proxrenart_VTI | bilateral | 125100\20354-7\G-A118\T-46600\T-71019 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL A PI | none | calc | c_Prox_SCL_A_PI | bilateral | 125100\12008-9\G-A118\T-45005\T-46100 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46100, Subclavian Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL A RI | none | calc | c_Prox_SCL_A_RI | bilateral | 125100\12023-8\G-A118\T-45005\T-46100 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46100, Subclavian Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SCL A S/D | none | calc | c_Prox_SCL_A_SD | bilateral | 125100\12144-2\G-A118\T-45005\T-46100 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46100, Subclavian Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFA PI | none | calc | c_Prox_SFA_PI | bilateral | 125100\12008-9\G-A118\T-47040\T-47403 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFA RI | none | calc | c_Prox_SFA_RI | bilateral | 125100\12023-8\G-A118\T-47040\T-47403 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SFA S/D | none | calc | c_Prox_SFA_SD | bilateral | 125100\12144-2\G-A118\T-47040\T-47403 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SMA AI | Doppler Slope | meas | m_proxsma_AS | unilateral | 125100\20167-3\G-A118\T-46002\T-46510 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SMA AT | Time | meas | m_proxsma_AT | unilateral | 125100\20168-1\G-A118\T-46002\T-46510 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SMA EDV | Signed End Diastolic Vel | meas | m_proxsma_EDV | unilateral | 125100\11653-3\G-A118\T-46002\T-46510 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SMA MDV | Signed Minimum Velocity | meas | m_proxsma_MDV | unilateral | 125100\11665-7\G-A118\T-46002\T-46510 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SMA PI | none | calc | c_proxsma_PI | unilateral | 125100\12008-9\G-A118\T-46002\T-46510 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SMA PSV | Signed Maximum Velocity | meas | m_proxsma_PSV | unilateral | 125100\11726-7\G-A118\T-46002\T-46510 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SMA RI | none | calc | c_proxsma_RI | unilateral | 125100\12023-8\G-A118\T-46002\T-46510 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SMA S/D | none | calc | c_proxsma_SD | unilateral | 125100\12144-2\G-A118\T-46002\T-46510 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SMA TAMV | Time Average Mean Velocity | meas | m_proxsma_TAMV | unilateral | 125100\20352-1\G-A118\T-46002\T-46510 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SMA TAPV | Signed Time Average Peak Vel | meas | m_proxsma_TAPV | unilateral | 125100\11692-1\G-A118\T-46002\T-46510 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox SMA VTI | Velocity Time Integral | meas | m_proxsma_VTI | unilateral | 125100\20354-7\G-A118\T-46002\T-46510 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46510, Superior Mesenteric Artery |  |  | SRT, T-46002, Artery of Abdomen | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar A PI | none | calc | c_Prox_Ulnar_A_PI | bilateral | 125100\12008-9\G-A118\T-47020\T-47200 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar A RI | none | calc | c_Prox_Ulnar_A_RI | bilateral | 125100\12023-8\G-A118\T-47020\T-47200 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | Prox Ulnar A S/D | none | calc | c_Prox_Ulnar_A_SD | bilateral | 125100\12144-2\G-A118\T-47020\T-47200 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Antecube_PSV | bilateral | 125100\11726-7\sup71_001\T-47020 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | 99PMSBLUS, sup71_001, Antecube |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Axillary_A_PSV | bilateral | 125100\11726-7\T-47020\T-47100 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47100, Axillary Artery |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Brach_Ceph_A_PSV | bilateral | 125100\11726-7\T-46010\T-47020 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46010, Innominate Artery |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Bulb_PSV | bilateral | 125100\11726-7\T-45005\T-45170 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45170, Carotid Bulb |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_CFA_PSV | bilateral | 125100\11726-7\T-47040\T-47400 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47400, Common Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Com_Iliac_A_PSV | bilateral | 125100\11726-7\T-46710\T-47040 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46710, Common Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Dist_ATA_PSV | bilateral | 125100\11726-7\G-A119\T-47040\T-47700 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Dist_Brachial_A_PSV | bilateral | 125100\11726-7\G-A119\T-47020\T-47160 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47160, Brachial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Dist_CCA_PSV | bilateral | 125100\11726-7\G-A119\T-45005\T-45100 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Dist_ICA_PSV | bilateral | 125100\11726-7\G-A119\T-45005\T-45300 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Dist_Pero_A_PSV | bilateral | 125100\11726-7\G-A119\T-47040\T-47630 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Dist_POPA_PSV | bilateral | 125100\11726-7\G-A119\T-47040\T-47500 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47500, Popliteal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Dist_PTA_PSV | bilateral | 125100\11726-7\G-A119\T-47040\T-47600 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Dist_Radial_A_PSV | bilateral | 125100\11726-7\G-A119\T-47020\T-47300 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Dist_SFA_PSV | bilateral | 125100\11726-7\G-A119\T-47040\T-47403 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Dist_Ulnar_A_PSV | bilateral | 125100\11726-7\G-A119\T-47020\T-47200 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Dor_Ped_ATA_PSV | bilateral | 125100\11726-7\T-47040\T-47741 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47741, Dorsalis Pedis Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Ext_Iliac_A_PSV | bilateral | 125100\11726-7\T-46910\T-47040 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46910, External Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Int_Iliac_A_PSV | bilateral | 125100\11726-7\T-46740\T-47040 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46740, Internal Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Mid_ATA_PSV | bilateral | 125100\11726-7\G-A188\T-47040\T-47700 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Mid_CCA_PSV | bilateral | 125100\11726-7\G-A188\T-45005\T-45100 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Mid_ECA_PSV | bilateral | 125100\11726-7\G-A188\T-45005\T-45200 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45200, External Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Mid_ICA_PSV | bilateral | 125100\11726-7\G-A188\T-45005\T-45300 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Mid_Pero_A_PSV | bilateral | 125100\11726-7\G-A188\T-47040\T-47630 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Mid_PTA_PSV | bilateral | 125100\11726-7\G-A188\T-47040\T-47600 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Mid_Radial_A_PSV | bilateral | 125100\11726-7\G-A188\T-47020\T-47300 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Mid_SCL_A_PSV | bilateral | 125100\11726-7\G-A188\T-45005\T-46100 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46100, Subclavian Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Mid_SFA_PSV | bilateral | 125100\11726-7\G-A188\T-47040\T-47403 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Mid_Ulnar_A_PSV | bilateral | 125100\11726-7\G-A188\T-47020\T-47200 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_PFA_PSV | bilateral | 125100\11726-7\G-A118\T-47040\T-47440 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47440, Profunda Femoris Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Prox_ATA_PSV | bilateral | 125100\11726-7\G-A118\T-47040\T-47700 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Prox_Brachial_A_PSV | bilateral | 125100\11726-7\G-A118\T-47020\T-47160 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47160, Brachial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Prox_CCA_PSV | bilateral | 125100\11726-7\G-A118\T-45005\T-45100 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Prox_ECA_PSV | bilateral | 125100\11726-7\G-A118\T-45005\T-45200 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45200, External Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Prox_ICA_PSV | bilateral | 125100\11726-7\G-A118\T-45005\T-45300 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Prox_Pero_A_PSV | bilateral | 125100\11726-7\G-A118\T-47040\T-47630 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Prox_POPA_PSV | bilateral | 125100\11726-7\G-A118\T-47040\T-47500 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47500, Popliteal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Prox_PTA_PSV | bilateral | 125100\11726-7\G-A118\T-47040\T-47600 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Prox_Radial_A_PSV | bilateral | 125100\11726-7\G-A118\T-47020\T-47300 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Prox_SCL_A_PSV | bilateral | 125100\11726-7\G-A118\T-45005\T-46100 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46100, Subclavian Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Prox_SFA_PSV | bilateral | 125100\11726-7\G-A118\T-47040\T-47403 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Prox_Ulnar_A_PSV | bilateral | 125100\11726-7\G-A118\T-47020\T-47200 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | PSV | Signed Maximum Velocity | meas | m_Vertebral_A_PSV | bilateral | 125100\11726-7\T-45005\T-45700 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45700, Vertebral Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | PV Bladder Height | Length | meas | m_pvbladder_ht | unilateral | 125100\C12011-07\C4-01\T5100-02 |  | 99PMSBLUS, C12011-07, Post Void Bladder Height |  | UCUM, cm, Centimeter | [6 empty] | 99PMSBLUS, C4-01, Post Void Bladder |  |  | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral)
Vascular Ultrasound Procedure Report | Vascular | PV Bladder Length | Length | meas | m_pvbladder_len | unilateral | 125100\C12011-05\C4-01\T5100-02 |  | 99PMSBLUS, C12011-05, Post Void Bladder Length |  | UCUM, cm, Centimeter | [6 empty] | 99PMSBLUS, C4-01, Post Void Bladder |  |  | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral)
Vascular Ultrasound Procedure Report | Vascular | PV Bladder Vol | none | calc | c_pvbladder_vol | unilateral | 125100\C12011-08\C4-01\T5100-02 |  | 99PMSBLUS, C12011-08, Post Void Bladder Volume |  | UCUM, cm3, Cubic Centimeter | [6 empty] | 99PMSBLUS, C4-01, Post Void Bladder |  |  | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral)
Vascular Ultrasound Procedure Report | Vascular | PV Bladder Width | Length | meas | m_pvbladder_wid | unilateral | 125100\C12011-06\C4-01\T5100-02 |  | 99PMSBLUS, C12011-06, Post Void Bladder Width |  | UCUM, cm, Centimeter | [6 empty] | 99PMSBLUS, C4-01, Post Void Bladder |  |  | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral)
Vascular Ultrasound Procedure Report | Vascular | R Hepatic A AI | Doppler Slope | meas | m_r_hepart_AS | unilateral | 125100\20167-3\T-46002\T-46423 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46423, Right Branch of Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | R Hepatic A AT | Time | meas | m_r_hepart_AT | unilateral | 125100\20168-1\T-46002\T-46423 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46423, Right Branch of Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | R Hepatic A EDV | Signed End Diastolic Vel | meas | m_r_hepart_EDV | unilateral | 125100\11653-3\T-46002\T-46423 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46423, Right Branch of Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | R Hepatic A MDV | Signed Minimum Velocity | meas | m_r_hepart_MDV | unilateral | 125100\11665-7\T-46002\T-46423 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46423, Right Branch of Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | R Hepatic A PI | none | calc | c_r_hepart_PI | unilateral | 125100\12008-9\T-46002\T-46423 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46423, Right Branch of Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | R Hepatic A PSV | Signed Maximum Velocity | meas | m_r_hepart_PSV | unilateral | 125100\11726-7\T-46002\T-46423 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46423, Right Branch of Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | R Hepatic A RI | none | calc | c_r_hepart_RI | unilateral | 125100\12023-8\T-46002\T-46423 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46423, Right Branch of Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | R Hepatic A S/D | none | calc | c_r_hepart_SD | unilateral | 125100\12144-2\T-46002\T-46423 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46423, Right Branch of Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | R Hepatic A TAMV | Time Average Mean Velocity | meas | m_r_hepart_TAMV | unilateral | 125100\20352-1\T-46002\T-46423 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46423, Right Branch of Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | R Hepatic A TAPV | Signed Time Average Peak Vel | meas | m_r_hepart_TAPV | unilateral | 125100\11692-1\T-46002\T-46423 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46423, Right Branch of Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | R Hepatic A VTI | Velocity Time Integral | meas | m_r_hepart_VTI | unilateral | 125100\20354-7\T-46002\T-46423 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46423, Right Branch of Hepatic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | RAR | none | calc | c_abdren_RAR | bilateral | 125100\33869-9\T-46600\T-71019 |  | LN, 33869-9, Renal Artery/Aorta velocity ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | RAR | none | calc | c_RAR | bilateral | 125100\33869-9\T-46600\T-71019 |  | LN, 33869-9, Renal Artery/Aorta velocity ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Ren A Org AI | Doppler Slope | meas | m_renaortorg_AS | bilateral | 125100\20167-3\G-036A\T-46600\T-71019 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Vascular | Ren A Org AT | Time | meas | m_renaortorg_AT | bilateral | 125100\20168-1\G-036A\T-46600\T-71019 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Vascular | Ren A Org EDV | Signed End Diastolic Vel | meas | m_renaortorg_EDV | bilateral | 125100\11653-3\G-036A\T-46600\T-71019 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Vascular | Ren A Org MDV | Signed Minimum Velocity | meas | m_renaortorg_MDV | bilateral | 125100\11665-7\G-036A\T-46600\T-71019 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Vascular | Ren A Org PI | none | calc | c_renaortorg_PI | bilateral | 125100\12008-9\G-036A\T-46600\T-71019 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Vascular | Ren A Org PSV | Signed Maximum Velocity | meas | m_renaortorg_PSV | bilateral | 125100\11726-7\G-036A\T-46600\T-71019 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Vascular | Ren A Org RI | none | calc | c_renaortorg_RI | bilateral | 125100\12023-8\G-036A\T-46600\T-71019 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Vascular | Ren A Org S/D | none | calc | c_renaortorg_SD | bilateral | 125100\12144-2\G-036A\T-46600\T-71019 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Vascular | Ren A Org TAMV | Time Average Mean Velocity | meas | m_renaortorg_TAMV | bilateral | 125100\20352-1\G-036A\T-46600\T-71019 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Vascular | Ren A Org TAPV | Signed Time Average Peak Vel | meas | m_renaortorg_TAPV | bilateral | 125100\11692-1\G-036A\T-46600\T-71019 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Vascular | Ren A Org VTI | Velocity Time Integral | meas | m_renaortorg_VTI | bilateral | 125100\20354-7\G-036A\T-46600\T-71019 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney | SRT, G-036A, Origin of vessel
Vascular Ultrasound Procedure Report | Vascular | Ren V AI | Doppler Slope | meas | m_renvein_AS | bilateral | 125100\20167-3\T-48740\T-71019 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-48740, Renal Vein |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Ren V AT | Time | meas | m_renvein_AT | bilateral | 125100\20168-1\T-48740\T-71019 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-48740, Renal Vein |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Ren V EDV | Signed End Diastolic Vel | meas | m_renvein_EDV | bilateral | 125100\11653-3\T-48740\T-71019 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-48740, Renal Vein |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Ren V MDV | Signed Minimum Velocity | meas | m_renvein_MDV | bilateral | 125100\11665-7\T-48740\T-71019 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-48740, Renal Vein |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Ren V PI | none | calc | c_renvein_PI | bilateral | 125100\12008-9\T-48740\T-71019 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-48740, Renal Vein |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Ren V PSV | Signed Maximum Velocity | meas | m_renvein_PSV | bilateral | 125100\11726-7\T-48740\T-71019 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-48740, Renal Vein |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Ren V RI | none | calc | c_renvein_RI | bilateral | 125100\12023-8\T-48740\T-71019 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-48740, Renal Vein |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Ren V S/D | none | calc | c_renvein_SD | bilateral | 125100\12144-2\T-48740\T-71019 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-48740, Renal Vein |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Ren V TAMV | Time Average Mean Velocity | meas | m_renvein_TAMV | bilateral | 125100\20352-1\T-48740\T-71019 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-48740, Renal Vein |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Ren V TAPV | Signed Time Average Peak Vel | meas | m_renvein_TAPV | bilateral | 125100\11692-1\T-48740\T-71019 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-48740, Renal Vein |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Ren V VTI | Velocity Time Integral | meas | m_renvein_VTI | bilateral | 125100\20354-7\T-48740\T-71019 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-48740, Renal Vein |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Renal A PSV | Signed Maximum Velocity | meas | m_abdren_renal_a_PSV | bilateral | 125100\11726-7\T-46600\T-71019 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Renal A PSV | Signed Maximum Velocity | meas | m_renal_a_PSV | bilateral | 125100\11726-7\T-46600\T-71019 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46600, Renal Artery |  |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | SFA % Area Reduction | none | calc | c_SFA_Area_Reduc | bilateral | 125100\G-0371\T-47040\T-47403 |  | SRT, G-0371, Percent Area Reduction |  | UCUM, %, Percent | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | SFA % Diam Reduction | none | calc | c_SFA_Diam_Reduc | bilateral | 125100\G-0372\T-47040\T-47403 |  | SRT, G-0372, Percent Diameter Reduction |  | UCUM, %, Percent | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | SFA Area Resid | Area | meas | m_SFA_Area_Resid | bilateral | 125100\C7471-01\T-47040\T-47403 |  | 99PMSBLUS, C7471-01, Area 1 of Area Percent Reduction |  | UCUM, cm2, Square Centimeter | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | SFA Area True | Area | meas | m_SFA_Area_True | bilateral | 125100\C7471-02\T-47040\T-47403 |  | 99PMSBLUS, C7471-02, Area 2 of Area Percent Reduction |  | UCUM, cm2, Square Centimeter | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | SFA Diam Resid | Length | meas | m_SFA_Diam_Resid | bilateral | 125100\C7470-05\T-47040\T-47403 |  | 99PMSBLUS, C7470-05, Diameter 1 of Diameter Reduction |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | SFA Diam True | Length | meas | m_SFA_Diam_True | bilateral | 125100\C7470-06\T-47040\T-47403 |  | 99PMSBLUS, C7470-06, Diameter 2 of Diameter Reduction |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | Spleen Height | Length | meas | m_splnHt | unilateral | 125100\C7470-01\T5100-02\T-C3000 |  | 99PMSBLUS, C7470-01, Height |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-C3000, Spleen |  |  | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral)
Vascular Ultrasound Procedure Report | Vascular | Spleen Length | Length | meas | m_splnLen | unilateral | 125100\G-A22A\T5100-02\T-C3000 |  | SRT, G-A22A, Length |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-C3000, Spleen |  |  | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral)
Vascular Ultrasound Procedure Report | Vascular | Spleen Vol | none | calc | c_splnvol | unilateral | 125100\121221\T5100-02\T-C3000 |  | DCM, 121221, Volume of ellipsoid |  | UCUM, cm3, Cubic Centimeter | [6 empty] | SRT, T-C3000, Spleen |  |  | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral)
Vascular Ultrasound Procedure Report | Vascular | Spleen Width | Length | meas | m_splnWid | unilateral | 125100\G-A220\T5100-02\T-C3000 |  | SRT, G-A220, Width |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-C3000, Spleen |  |  | 99PMSBLUS, T5100-02, Anatomic Structures (unilateral)
Vascular Ultrasound Procedure Report | Vascular | Splenic A AI | Doppler Slope | meas | m_splenicart_AS | unilateral | 125100\20167-3\T-46002\T-46460 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46460, Splenic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Splenic A AT | Time | meas | m_splenicart_AT | unilateral | 125100\20168-1\T-46002\T-46460 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46460, Splenic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Splenic A EDV | Signed End Diastolic Vel | meas | m_splenicart_EDV | unilateral | 125100\11653-3\T-46002\T-46460 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46460, Splenic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Splenic A MDV | Signed Minimum Velocity | meas | m_splenicart_MDV | unilateral | 125100\11665-7\T-46002\T-46460 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46460, Splenic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Splenic A PI | none | calc | c_splenicart_PI | unilateral | 125100\12008-9\T-46002\T-46460 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46460, Splenic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Splenic A PSV | Signed Maximum Velocity | meas | m_splenicart_PSV | unilateral | 125100\11726-7\T-46002\T-46460 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46460, Splenic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Splenic A RI | none | calc | c_splenicart_RI | unilateral | 125100\12023-8\T-46002\T-46460 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46460, Splenic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Splenic A S/D | none | calc | c_splenicart_SD | unilateral | 125100\12144-2\T-46002\T-46460 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46460, Splenic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Splenic A TAMV | Time Average Mean Velocity | meas | m_splenicart_TAMV | unilateral | 125100\20352-1\T-46002\T-46460 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46460, Splenic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Splenic A TAPV | Signed Time Average Peak Vel | meas | m_splenicart_TAPV | unilateral | 125100\11692-1\T-46002\T-46460 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46460, Splenic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Splenic A VTI | Velocity Time Integral | meas | m_splenicart_VTI | unilateral | 125100\20354-7\T-46002\T-46460 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46460, Splenic Artery |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Sup Arc AI | Doppler Slope | meas | m_suparc_AS | bilateral | 125100\20167-3\G-A116\T-4668A\T-71019 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Sup Arc AT | Time | meas | m_suparc_AT | bilateral | 125100\20168-1\G-A116\T-4668A\T-71019 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Sup Arc EDV | Signed End Diastolic Vel | meas | m_suparc_EDV | bilateral | 125100\11653-3\G-A116\T-4668A\T-71019 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Sup Arc MDV | Signed Minimum Velocity | meas | m_suparc_MDV | bilateral | 125100\11665-7\G-A116\T-4668A\T-71019 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Sup Arc PI | none | calc | c_suparc_PI | bilateral | 125100\12008-9\G-A116\T-4668A\T-71019 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Sup Arc PSV | Signed Maximum Velocity | meas | m_suparc_PSV | bilateral | 125100\11726-7\G-A116\T-4668A\T-71019 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Sup Arc RI | none | calc | c_suparc_RI | bilateral | 125100\12023-8\G-A116\T-4668A\T-71019 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Sup Arc S/D | none | calc | c_suparc_SD | bilateral | 125100\12144-2\G-A116\T-4668A\T-71019 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Sup Arc TAMV | Time Average Mean Velocity | meas | m_suparc_TAMV | bilateral | 125100\20352-1\G-A116\T-4668A\T-71019 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Sup Arc TAPV | Signed Time Average Peak Vel | meas | m_suparc_TAPV | bilateral | 125100\11692-1\G-A116\T-4668A\T-71019 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Sup Arc VTI | Velocity Time Integral | meas | m_suparc_VTI | bilateral | 125100\20354-7\G-A116\T-4668A\T-71019 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-4668A, Arcuate Artery of the Kidney | SRT, G-A116, Superior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Sup Ren Ao AI | Doppler Slope | meas | m_suprenaorta_AS | unilateral | 125100\20167-3\T-42510\T-46002 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-42510, Supra-renal Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Sup Ren Ao AT | Time | meas | m_suprenaorta_AT | unilateral | 125100\20168-1\T-42510\T-46002 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-42510, Supra-renal Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Sup Ren Ao EDV | Signed End Diastolic Vel | meas | m_suprenaorta_EDV | unilateral | 125100\11653-3\T-42510\T-46002 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-42510, Supra-renal Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Sup Ren Ao MDV | Signed Minimum Velocity | meas | m_suprenaorta_MDV | unilateral | 125100\11665-7\T-42510\T-46002 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-42510, Supra-renal Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Sup Ren Ao PI | none | calc | c_suprenaorta_PI | unilateral | 125100\12008-9\T-42510\T-46002 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-42510, Supra-renal Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Sup Ren Ao PSV | Signed Maximum Velocity | meas | m_suprenaorta_PSV | unilateral | 125100\11726-7\T-42510\T-46002 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-42510, Supra-renal Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Sup Ren Ao RI | none | calc | c_suprenaorta_RI | unilateral | 125100\12023-8\T-42510\T-46002 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-42510, Supra-renal Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Sup Ren Ao S/D | none | calc | c_suprenaorta_SD | unilateral | 125100\12144-2\T-42510\T-46002 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-42510, Supra-renal Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Sup Ren Ao TAMV | Time Average Mean Velocity | meas | m_suprenaorta_TAMV | unilateral | 125100\20352-1\T-42510\T-46002 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-42510, Supra-renal Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Sup Ren Ao TAPV | Signed Time Average Peak Vel | meas | m_suprenaorta_TAPV | unilateral | 125100\11692-1\T-42510\T-46002 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-42510, Supra-renal Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Sup Ren Ao VTI | Velocity Time Integral | meas | m_suprenaorta_VTI | unilateral | 125100\20354-7\T-42510\T-46002 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-42510, Supra-renal Aorta |  |  | SRT, T-46002, Artery of Abdomen
Vascular Ultrasound Procedure Report | Vascular | Sup Seg AI | Doppler Slope | meas | m_supseg_AS | bilateral | 125100\20167-3\G-A116\T-46659\T-71019 |  | LN, 20167-3, Acceleration Index |  | UCUM, cm/s2, Centimeter Per Second Square | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Sup Seg AT | Time | meas | m_supseg_AT | bilateral | 125100\20168-1\G-A116\T-46659\T-71019 |  | LN, 20168-1, Acceleration Time |  | UCUM, sec, Seconds | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Sup Seg EDV | Signed End Diastolic Vel | meas | m_supseg_EDV | bilateral | 125100\11653-3\G-A116\T-46659\T-71019 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Sup Seg MDV | Signed Minimum Velocity | meas | m_supseg_MDV | bilateral | 125100\11665-7\G-A116\T-46659\T-71019 |  | LN, 11665-7, Minimum Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Sup Seg PI | none | calc | c_supseg_PI | bilateral | 125100\12008-9\G-A116\T-46659\T-71019 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Sup Seg PSV | Signed Maximum Velocity | meas | m_supseg_PSV | bilateral | 125100\11726-7\G-A116\T-46659\T-71019 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Sup Seg RI | none | calc | c_supseg_RI | bilateral | 125100\12023-8\G-A116\T-46659\T-71019 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Sup Seg S/D | none | calc | c_supseg_SD | bilateral | 125100\12144-2\G-A116\T-46659\T-71019 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Sup Seg TAMV | Time Average Mean Velocity | meas | m_supseg_TAMV | bilateral | 125100\20352-1\G-A116\T-46659\T-71019 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Sup Seg TAPV | Signed Time Average Peak Vel | meas | m_supseg_TAPV | bilateral | 125100\11692-1\G-A116\T-46659\T-71019 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | Sup Seg VTI | Velocity Time Integral | meas | m_supseg_VTI | bilateral | 125100\20354-7\G-A116\T-46659\T-71019 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46659, Segmental Artery | SRT, G-A116, Superior |  | SRT, T-71019, Vascular Structure Of Kidney
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Antecube_TAMV | bilateral | 125100\20352-1\sup71_001\T-47020 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | 99PMSBLUS, sup71_001, Antecube |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Axillary_A_TAMV | bilateral | 125100\20352-1\T-47020\T-47100 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47100, Axillary Artery |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Brach_Ceph_A_TAMV | bilateral | 125100\20352-1\T-46010\T-47020 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46010, Innominate Artery |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Bulb_TAMV | bilateral | 125100\20352-1\T-45005\T-45170 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45170, Carotid Bulb |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_CFA_TAMV | bilateral | 125100\20352-1\T-47040\T-47400 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47400, Common Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Com_Iliac_A_TAMV | bilateral | 125100\20352-1\T-46710\T-47040 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46710, Common Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Dist_ATA_TAMV | bilateral | 125100\20352-1\G-A119\T-47040\T-47700 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Dist_Brachial_A_TAMV | bilateral | 125100\20352-1\G-A119\T-47020\T-47160 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47160, Brachial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Dist_CCA_TAMV | bilateral | 125100\20352-1\G-A119\T-45005\T-45100 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Dist_ICA_TAMV | bilateral | 125100\20352-1\G-A119\T-45005\T-45300 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Dist_Pero_A_TAMV | bilateral | 125100\20352-1\G-A119\T-47040\T-47630 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Dist_POPA_TAMV | bilateral | 125100\20352-1\G-A119\T-47040\T-47500 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47500, Popliteal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Dist_PTA_TAMV | bilateral | 125100\20352-1\G-A119\T-47040\T-47600 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Dist_Radial_A_TAMV | bilateral | 125100\20352-1\G-A119\T-47020\T-47300 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Dist_SFA_TAMV | bilateral | 125100\20352-1\G-A119\T-47040\T-47403 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Dist_Ulnar_A_TAMV | bilateral | 125100\20352-1\G-A119\T-47020\T-47200 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Dor_Ped_ATA_TAMV | bilateral | 125100\20352-1\T-47040\T-47741 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47741, Dorsalis Pedis Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Ext_Iliac_A_TAMV | bilateral | 125100\20352-1\T-46910\T-47040 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46910, External Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Int_Iliac_A_TAMV | bilateral | 125100\20352-1\T-46740\T-47040 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46740, Internal Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Mid_ATA_TAMV | bilateral | 125100\20352-1\G-A188\T-47040\T-47700 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Mid_CCA_TAMV | bilateral | 125100\20352-1\G-A188\T-45005\T-45100 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Mid_ECA_TAMV | bilateral | 125100\20352-1\G-A188\T-45005\T-45200 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45200, External Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Mid_ICA_TAMV | bilateral | 125100\20352-1\G-A188\T-45005\T-45300 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Mid_Pero_A_TAMV | bilateral | 125100\20352-1\G-A188\T-47040\T-47630 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Mid_PTA_TAMV | bilateral | 125100\20352-1\G-A188\T-47040\T-47600 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Mid_Radial_A_TAMV | bilateral | 125100\20352-1\G-A188\T-47020\T-47300 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Mid_SCL_A_TAMV | bilateral | 125100\20352-1\G-A188\T-45005\T-46100 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46100, Subclavian Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Mid_SFA_TAMV | bilateral | 125100\20352-1\G-A188\T-47040\T-47403 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Mid_Ulnar_A_TAMV | bilateral | 125100\20352-1\G-A188\T-47020\T-47200 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_PFA_TAMV | bilateral | 125100\20352-1\G-A118\T-47040\T-47440 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47440, Profunda Femoris Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Prox_ATA_TAMV | bilateral | 125100\20352-1\G-A118\T-47040\T-47700 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Prox_Brachial_A_TAMV | bilateral | 125100\20352-1\G-A118\T-47020\T-47160 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47160, Brachial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Prox_CCA_TAMV | bilateral | 125100\20352-1\G-A118\T-45005\T-45100 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Prox_ECA_TAMV | bilateral | 125100\20352-1\G-A118\T-45005\T-45200 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45200, External Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Prox_ICA_TAMV | bilateral | 125100\20352-1\G-A118\T-45005\T-45300 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Prox_Pero_A_TAMV | bilateral | 125100\20352-1\G-A118\T-47040\T-47630 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Prox_POPA_TAMV | bilateral | 125100\20352-1\G-A118\T-47040\T-47500 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47500, Popliteal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Prox_PTA_TAMV | bilateral | 125100\20352-1\G-A118\T-47040\T-47600 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Prox_Radial_A_TAMV | bilateral | 125100\20352-1\G-A118\T-47020\T-47300 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Prox_SCL_A_TAMV | bilateral | 125100\20352-1\G-A118\T-45005\T-46100 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46100, Subclavian Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Prox_SFA_TAMV | bilateral | 125100\20352-1\G-A118\T-47040\T-47403 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Prox_Ulnar_A_TAMV | bilateral | 125100\20352-1\G-A118\T-47020\T-47200 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAMV | Time Average Mean Velocity | meas | m_Vertebral_A_TAMV | bilateral | 125100\20352-1\T-45005\T-45700 |  | LN, 20352-1, Time averaged mean velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45700, Vertebral Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Antecube_TAPV | bilateral | 125100\11692-1\sup71_001\T-47020 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | 99PMSBLUS, sup71_001, Antecube |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Axillary_A_TAPV | bilateral | 125100\11692-1\T-47020\T-47100 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47100, Axillary Artery |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Brach_Ceph_A_TAPV | bilateral | 125100\11692-1\T-46010\T-47020 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46010, Innominate Artery |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Bulb_TAPV | bilateral | 125100\11692-1\T-45005\T-45170 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45170, Carotid Bulb |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_CFA_TAPV | bilateral | 125100\11692-1\T-47040\T-47400 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47400, Common Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Com_Iliac_A_TAPV | bilateral | 125100\11692-1\T-46710\T-47040 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46710, Common Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Dist_ATA_TAPV | bilateral | 125100\11692-1\G-A119\T-47040\T-47700 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Dist_Brachial_A_TAPV | bilateral | 125100\11692-1\G-A119\T-47020\T-47160 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47160, Brachial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Dist_CCA_TAPV | bilateral | 125100\11692-1\G-A119\T-45005\T-45100 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Dist_ICA_TAPV | bilateral | 125100\11692-1\G-A119\T-45005\T-45300 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Dist_Pero_A_TAPV | bilateral | 125100\11692-1\G-A119\T-47040\T-47630 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Dist_POPA_TAPV | bilateral | 125100\11692-1\G-A119\T-47040\T-47500 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47500, Popliteal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Dist_PTA_TAPV | bilateral | 125100\11692-1\G-A119\T-47040\T-47600 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Dist_Radial_A_TAPV | bilateral | 125100\11692-1\G-A119\T-47020\T-47300 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Dist_SFA_TAPV | bilateral | 125100\11692-1\G-A119\T-47040\T-47403 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Dist_Ulnar_A_TAPV | bilateral | 125100\11692-1\G-A119\T-47020\T-47200 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Dor_Ped_ATA_TAPV | bilateral | 125100\11692-1\T-47040\T-47741 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47741, Dorsalis Pedis Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Ext_Iliac_A_TAPV | bilateral | 125100\11692-1\T-46910\T-47040 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46910, External Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Int_Iliac_A_TAPV | bilateral | 125100\11692-1\T-46740\T-47040 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46740, Internal Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Mid_ATA_TAPV | bilateral | 125100\11692-1\G-A188\T-47040\T-47700 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Mid_CCA_TAPV | bilateral | 125100\11692-1\G-A188\T-45005\T-45100 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Mid_ECA_TAPV | bilateral | 125100\11692-1\G-A188\T-45005\T-45200 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45200, External Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Mid_ICA_TAPV | bilateral | 125100\11692-1\G-A188\T-45005\T-45300 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Mid_Pero_A_TAPV | bilateral | 125100\11692-1\G-A188\T-47040\T-47630 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Mid_PTA_TAPV | bilateral | 125100\11692-1\G-A188\T-47040\T-47600 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Mid_Radial_A_TAPV | bilateral | 125100\11692-1\G-A188\T-47020\T-47300 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Mid_SCL_A_TAPV | bilateral | 125100\11692-1\G-A188\T-45005\T-46100 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46100, Subclavian Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Mid_SFA_TAPV | bilateral | 125100\11692-1\G-A188\T-47040\T-47403 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Mid_Ulnar_A_TAPV | bilateral | 125100\11692-1\G-A188\T-47020\T-47200 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_PFA_TAPV | bilateral | 125100\11692-1\G-A118\T-47040\T-47440 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47440, Profunda Femoris Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Prox_ATA_TAPV | bilateral | 125100\11692-1\G-A118\T-47040\T-47700 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Prox_Brachial_A_TAPV | bilateral | 125100\11692-1\G-A118\T-47020\T-47160 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47160, Brachial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Prox_CCA_TAPV | bilateral | 125100\11692-1\G-A118\T-45005\T-45100 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Prox_ECA_TAPV | bilateral | 125100\11692-1\G-A118\T-45005\T-45200 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45200, External Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Prox_ICA_TAPV | bilateral | 125100\11692-1\G-A118\T-45005\T-45300 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Prox_Pero_A_TAPV | bilateral | 125100\11692-1\G-A118\T-47040\T-47630 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Prox_POPA_TAPV | bilateral | 125100\11692-1\G-A118\T-47040\T-47500 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47500, Popliteal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Prox_PTA_TAPV | bilateral | 125100\11692-1\G-A118\T-47040\T-47600 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Prox_Radial_A_TAPV | bilateral | 125100\11692-1\G-A118\T-47020\T-47300 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Prox_SCL_A_TAPV | bilateral | 125100\11692-1\G-A118\T-45005\T-46100 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-46100, Subclavian Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Prox_SFA_TAPV | bilateral | 125100\11692-1\G-A118\T-47040\T-47403 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Prox_Ulnar_A_TAPV | bilateral | 125100\11692-1\G-A118\T-47020\T-47200 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | TAPV | Signed Time Average Peak Vel | meas | m_Vertebral_A_TAPV | bilateral | 125100\11692-1\T-45005\T-45700 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, T-45700, Vertebral Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | Term ICA EDV | Signed End Diastolic Vel | meas | m_tcd_Term_ICA_EDV | bilateral | 125100\11653-3\R-102BD\T-40501 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-102BD, Terminal internal carotid artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Term ICA PI | none | calc | c_tcd_Term_ICA_PI | bilateral | 125100\12008-9\R-102BD\T-40501 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, R-102BD, Terminal internal carotid artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Term ICA PSV | Signed Maximum Velocity | meas | m_tcd_Term_ICA_PSV | bilateral | 125100\11726-7\R-102BD\T-40501 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-102BD, Terminal internal carotid artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Term ICA TAPV | Signed Time Average Peak Vel | meas | m_tcd_Term_ICA_TAPV | bilateral | 125100\11692-1\R-102BD\T-40501 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-102BD, Terminal internal carotid artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Term ICA TCD Mean | none | calc | c_tcd_Term_ICA_TcdMean | bilateral | 125100\C12221-03\R-102BD\T-40501 |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | SRT, R-102BD, Terminal internal carotid artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Term ICA TCD PI | none | calc | c_tcd_Term_ICA_TcdPI | bilateral | 125100\C12221-04\R-102BD\T-40501 |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, R-102BD, Terminal internal carotid artery |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Vertebral A EDV | Signed End Diastolic Vel | meas | m_tcd_Vertebral_A_EDV | bilateral | 125100\11653-3\C12105-04\T-40501 |  | LN, 11653-3, End Diastolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Vertebral A PI | none | calc | c_tcd_Vertebral_A_PI | bilateral | 125100\12008-9\C12105-04\T-40501 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Vertebral A PI | none | calc | c_Vertebral_A_PI | bilateral | 125100\12008-9\T-45005\T-45700 |  | LN, 12008-9, Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45700, Vertebral Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | Vertebral A PSV | Signed Maximum Velocity | meas | m_tcd_Vertebral_A_PSV | bilateral | 125100\11726-7\C12105-04\T-40501 |  | LN, 11726-7, Peak Systolic Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Vertebral A RI | none | calc | c_Vertebral_A_RI | bilateral | 125100\12023-8\T-45005\T-45700 |  | LN, 12023-8, Resistivity Index |  | UCUM, 1, ratio | [6 empty] | SRT, T-45700, Vertebral Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | Vertebral A S/D | none | calc | c_Vertebral_A_SD | bilateral | 125100\12144-2\T-45005\T-45700 |  | LN, 12144-2, Systolic to Diastolic Velocity Ratio |  | UCUM, 1, ratio | [6 empty] | SRT, T-45700, Vertebral Artery |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | Vertebral A TAPV | Signed Time Average Peak Vel | meas | m_tcd_Vertebral_A_TAPV | bilateral | 125100\11692-1\C12105-04\T-40501 |  | LN, 11692-1, Time averaged peak velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Vertebral A TCD Mean | none | calc | c_tcd_Vertebral_A_TcdMean | bilateral | 125100\C12105-04\C12221-03\T-40501 |  | 99PMSBLUS, C12221-03, TCD Mean Velocity |  | UCUM, cm/s, Centimeter Per Second | [6 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | Vertebral A TCD PI | none | calc | c_tcd_Vertebral_A_TcdPI | bilateral | 125100\C12105-04\C12221-04\T-40501 |  | 99PMSBLUS, C12221-04, TCD Pulsatility Index |  | UCUM, 1, ratio | [6 empty] | 99PMSBLUS, C12105-04, Vertebral Artery in TCD |  |  | SRT, T-40501, Blood Vessel of Head
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Antecube_VTI | bilateral | 125100\20354-7\sup71_001\T-47020 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | 99PMSBLUS, sup71_001, Antecube |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Axillary_A_VTI | bilateral | 125100\20354-7\T-47020\T-47100 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47100, Axillary Artery |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Brach_Ceph_A_VTI | bilateral | 125100\20354-7\T-46010\T-47020 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46010, Innominate Artery |  |  | SRT, T-47020, Artery Of Upper Extremity
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Bulb_VTI | bilateral | 125100\20354-7\T-45005\T-45170 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45170, Carotid Bulb |  |  | SRT, T-45005, Artery of neck
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_CFA_VTI | bilateral | 125100\20354-7\T-47040\T-47400 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47400, Common Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Com_Iliac_A_VTI | bilateral | 125100\20354-7\T-46710\T-47040 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46710, Common Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Dist_ATA_VTI | bilateral | 125100\20354-7\G-A119\T-47040\T-47700 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Dist_Brachial_A_VTI | bilateral | 125100\20354-7\G-A119\T-47020\T-47160 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47160, Brachial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Dist_CCA_VTI | bilateral | 125100\20354-7\G-A119\T-45005\T-45100 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Dist_ICA_VTI | bilateral | 125100\20354-7\G-A119\T-45005\T-45300 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Dist_Pero_A_VTI | bilateral | 125100\20354-7\G-A119\T-47040\T-47630 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Dist_POPA_VTI | bilateral | 125100\20354-7\G-A119\T-47040\T-47500 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47500, Popliteal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Dist_PTA_VTI | bilateral | 125100\20354-7\G-A119\T-47040\T-47600 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Dist_Radial_A_VTI | bilateral | 125100\20354-7\G-A119\T-47020\T-47300 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Dist_SFA_VTI | bilateral | 125100\20354-7\G-A119\T-47040\T-47403 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Dist_Ulnar_A_VTI | bilateral | 125100\20354-7\G-A119\T-47020\T-47200 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A119, Distal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Dor_Ped_ATA_VTI | bilateral | 125100\20354-7\T-47040\T-47741 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47741, Dorsalis Pedis Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Ext_Iliac_A_VTI | bilateral | 125100\20354-7\T-46910\T-47040 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46910, External Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Int_Iliac_A_VTI | bilateral | 125100\20354-7\T-46740\T-47040 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46740, Internal Iliac Artery |  |  | SRT, T-47040, Artery of Lower Extremity
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Mid_ATA_VTI | bilateral | 125100\20354-7\G-A188\T-47040\T-47700 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Mid_CCA_VTI | bilateral | 125100\20354-7\G-A188\T-45005\T-45100 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Mid_ECA_VTI | bilateral | 125100\20354-7\G-A188\T-45005\T-45200 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45200, External Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Mid_ICA_VTI | bilateral | 125100\20354-7\G-A188\T-45005\T-45300 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Mid_Pero_A_VTI | bilateral | 125100\20354-7\G-A188\T-47040\T-47630 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Mid_PTA_VTI | bilateral | 125100\20354-7\G-A188\T-47040\T-47600 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Mid_Radial_A_VTI | bilateral | 125100\20354-7\G-A188\T-47020\T-47300 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Mid_SCL_A_VTI | bilateral | 125100\20354-7\G-A188\T-45005\T-46100 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46100, Subclavian Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Mid_SFA_VTI | bilateral | 125100\20354-7\G-A188\T-47040\T-47403 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Mid_Ulnar_A_VTI | bilateral | 125100\20354-7\G-A188\T-47020\T-47200 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A188, Mid-longitudinal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_PFA_VTI | bilateral | 125100\20354-7\G-A118\T-47040\T-47440 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47440, Profunda Femoris Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Prox_ATA_VTI | bilateral | 125100\20354-7\G-A118\T-47040\T-47700 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47700, Anterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Prox_Brachial_A_VTI | bilateral | 125100\20354-7\G-A118\T-47020\T-47160 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47160, Brachial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Prox_CCA_VTI | bilateral | 125100\20354-7\G-A118\T-45005\T-45100 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45100, Common Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Prox_ECA_VTI | bilateral | 125100\20354-7\G-A118\T-45005\T-45200 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45200, External Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Prox_ICA_VTI | bilateral | 125100\20354-7\G-A118\T-45005\T-45300 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45300, Internal Carotid Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Prox_Pero_A_VTI | bilateral | 125100\20354-7\G-A118\T-47040\T-47630 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47630, Peroneal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Prox_POPA_VTI | bilateral | 125100\20354-7\G-A118\T-47040\T-47500 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47500, Popliteal Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Prox_PTA_VTI | bilateral | 125100\20354-7\G-A118\T-47040\T-47600 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47600, Posterior Tibial Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Prox_Radial_A_VTI | bilateral | 125100\20354-7\G-A118\T-47020\T-47300 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47300, Radial Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Prox_SCL_A_VTI | bilateral | 125100\20354-7\G-A118\T-45005\T-46100 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-46100, Subclavian Artery |  |  | SRT, T-45005, Artery of neck | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Prox_SFA_VTI | bilateral | 125100\20354-7\G-A118\T-47040\T-47403 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47403, Superficial Femoral Artery |  |  | SRT, T-47040, Artery of Lower Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Prox_Ulnar_A_VTI | bilateral | 125100\20354-7\G-A118\T-47020\T-47200 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-47200, Ulnar Artery |  |  | SRT, T-47020, Artery Of Upper Extremity | SRT, G-A118, Proximal
Vascular Ultrasound Procedure Report | Vascular | VTI | Velocity Time Integral | meas | m_Vertebral_A_VTI | bilateral | 125100\20354-7\T-45005\T-45700 |  | LN, 20354-7, Velocity Time Integral |  | UCUM, cm, Centimeter | [6 empty] | SRT, T-45700, Vertebral Artery |  |  | SRT, T-45005, Artery of neck